## VNA
Impedance-measurement errors with VNA S11 and S21 data
| C | R | L | Scaling error | Noise floor | Quantization
| 4.7e-05 | 0.005 | 5e-10 | 0.999 | 1e-12 | 1e-12
Press F9 to update noise!
N_sweep | Offset_error | F_min | F_max
500 | 0 | 1000 | 1000000000
N_steps | incr
201 | 1.0715
|  | ****************** D U T ********************* | [6 empty] | ADD VNA ERROR |  | ADD NOISE |  | QUANTIZATION ERROR |  |  |  | CALCULATE IMPEDANCE | [6 empty] | ADD VNA ERROR |  | ADD NOISE |  | QUANTIZATION ERROR |  |  |  | CALCULATE IMPEDANCE
n | Freq | Zreal | Zimag | Mag_Z | Phase |  | Re_S21 | Im_S21 | Re_S21 | Im_S21 | Re_S21 | Im_S21 | Re_S21 | Im_S21 | S21_Mag | S21_Phase | Re_Z | Im_Z | Mag_Z | Phase | Mag_err | Re_S11 | Im_S11 | Re_S11 | Im_S11 | Re_S11 | Im_S11 | Re_S11 | Im_S11 | S11_Mag | S11_Phase | Re_Z | Im_Z | Mag_Z | Phase | Mag_err
1 | 1000 | 0.005 | -3.3863 | 3.3863 | -89.9154 |  | 0.018206 | -0.13296 | 0.018187 | -0.13283 | 0.018187 | -0.13283 | 0.018187 | -0.13283 | -17.4537 | -82.2031 | 0.0054532 | -3.3829 | 3.3829 | -89.9076 | -0.10002 | -0.99067 | -0.13481 | -0.98968 | -0.13467 | -0.98968 | -0.13467 | -0.98968 | -0.13467 | -0.010419 | -172.2511 | 0.030127 | -3.3863 | 3.3864 | -89.4903 | 0.0038135
2 | 1071.5193 | 0.005 | -3.1603 | 3.1603 | -89.9093 |  | 0.015919 | -0.12437 | 0.015903 | -0.12425 | 0.015903 | -0.12425 | 0.015903 | -0.12425 | -18.0436 | -82.7062 | 0.0053941 | -3.1571 | 3.1571 | -89.9021 | -0.10002 | -0.99184 | -0.12588 | -0.99085 | -0.12576 | -0.99085 | -0.12576 | -0.99085 | -0.12576 | -0.010421 | -172.7669 | 0.030112 | -3.1603 | 3.1604 | -89.4541 | 0.0043792
3 | 1148.1536 | 0.005 | -2.9493 | 2.9493 | -89.9029 |  | 0.013918 | -0.11631 | 0.013904 | -0.11619 | 0.013904 | -0.11619 | 0.013904 | -0.11619 | -18.6348 | -83.176 | 0.0053426 | -2.9464 | 2.9464 | -89.8961 | -0.10002 | -0.99287 | -0.11754 | -0.99187 | -0.11742 | -0.99187 | -0.11742 | -0.99187 | -0.11742 | -0.010421 | -173.2485 | 0.0301 | -2.9493 | 2.9495 | -89.4153 | 0.0050287
4 | 1230.2688 | 0.005 | -2.7525 | 2.7525 | -89.8959 |  | 0.012169 | -0.10874 | 0.012157 | -0.10863 | 0.012157 | -0.10863 | 0.012157 | -0.10863 | -19.2271 | -83.6143 | 0.0052977 | -2.7497 | 2.7497 | -89.8896 | -0.10002 | -0.99376 | -0.10974 | -0.99277 | -0.10963 | -0.99277 | -0.10963 | -0.99277 | -0.10963 | -0.010422 | -173.6982 | 0.030088 | -2.7525 | 2.7526 | -89.3737 | 0.0057745
5 | 1318.2567 | 0.005 | -2.5687 | 2.5688 | -89.8885 |  | 0.010641 | -0.10164 | 0.01063 | -0.10153 | 0.01063 | -0.10153 | 0.01063 | -0.10153 | -19.8204 | -84.0231 | 0.0052587 | -2.5662 | 2.5662 | -89.8826 | -0.10002 | -0.99454 | -0.10246 | -0.99354 | -0.10236 | -0.99354 | -0.10236 | -0.99354 | -0.10236 | -0.010423 | -174.118 | 0.030079 | -2.5687 | 2.5689 | -89.3291 | 0.0066307
6 | 1412.5375 | 0.005 | -2.3973 | 2.3973 | -89.8805 |  | 0.009306 | -0.09498 | 0.0092967 | -0.094885 | 0.0092967 | -0.094885 | 0.0092967 | -0.094885 | -20.4145 | -84.4041 | 0.0052246 | -2.3949 | 2.3949 | -89.875 | -0.10002 | -0.99521 | -0.095653 | -0.99422 | -0.095557 | -0.99422 | -0.095557 | -0.99422 | -0.095557 | -0.010423 | -174.51 | 0.03007 | -2.3973 | 2.3975 | -89.2814 | 0.0076138
7 | 1513.5612 | 0.005 | -2.2373 | 2.2373 | -89.872 |  | 0.0081403 | -0.088745 | 0.0081322 | -0.088656 | 0.0081322 | -0.088656 | 0.0081322 | -0.088656 | -21.0094 | -84.7591 | 0.005195 | -2.235 | 2.2351 | -89.8668 | -0.10002 | -0.9958 | -0.089295 | -0.99481 | -0.089205 | -0.99481 | -0.089205 | -0.99481 | -0.089205 | -0.010424 | -174.8759 | 0.030063 | -2.2373 | 2.2375 | -89.2302 | 0.0087425
8 | 1621.8101 | 0.005 | -2.088 | 2.088 | -89.8628 |  | 0.0071228 | -0.082907 | 0.0071157 | -0.082824 | 0.0071157 | -0.082824 | 0.0071157 | -0.082824 | -21.605 | -85.0896 | 0.0051692 | -2.0859 | 2.0859 | -89.858 | -0.10002 | -0.99632 | -0.083356 | -0.99532 | -0.083273 | -0.99532 | -0.083273 | -0.99532 | -0.083273 | -0.010424 | -175.2175 | 0.030056 | -2.088 | 2.0882 | -89.1753 | 0.010038
9 | 1737.8008 | 0.005 | -1.9486 | 1.9486 | -89.853 |  | 0.0062349 | -0.077442 | 0.0062287 | -0.077365 | 0.0062287 | -0.077365 | 0.0062287 | -0.077365 | -22.2011 | -85.397 | 0.0051467 | -1.9466 | 1.9467 | -89.8485 | -0.10002 | -0.99677 | -0.07781 | -0.99577 | -0.077732 | -0.99577 | -0.077732 | -0.99577 | -0.077732 | -0.010425 | -175.5364 | 0.03005 | -1.9486 | 1.9488 | -89.1165 | 0.011526
10 | 1862.0871 | 0.005 | -1.8185 | 1.8185 | -89.8425 |  | 0.0054603 | -0.07233 | 0.0054548 | -0.072257 | 0.0054548 | -0.072257 | 0.0054548 | -0.072257 | -22.7977 | -85.6829 | 0.0051271 | -1.8167 | 1.8167 | -89.8383 | -0.10002 | -0.99716 | -0.072631 | -0.99616 | -0.072558 | -0.99616 | -0.072558 | -0.99616 | -0.072558 | -0.010425 | -175.8341 | 0.030046 | -1.8185 | 1.8188 | -89.0535 | 0.013235
11 | 1995.2623 | 0.005 | -1.6972 | 1.6972 | -89.8312 |  | 0.0047846 | -0.067548 | 0.0047798 | -0.06748 | 0.0047798 | -0.06748 | 0.0047798 | -0.06748 | -23.3947 | -85.9483 | 0.0051101 | -1.6955 | 1.6955 | -89.8273 | -0.10002 | -0.9975 | -0.067794 | -0.9965 | -0.067727 | -0.9965 | -0.067727 | -0.9965 | -0.067727 | -0.010425 | -176.1119 | 0.030041 | -1.6972 | 1.6974 | -88.9859 | 0.015196
12 | 2137.9621 | 0.005 | -1.5839 | 1.5839 | -89.8191 |  | 0.0041954 | -0.063077 | 0.0041912 | -0.063013 | 0.0041912 | -0.063013 | 0.0041912 | -0.063013 | -23.9922 | -86.1947 | 0.0050952 | -1.5823 | 1.5823 | -89.8155 | -0.10002 | -0.9978 | -0.063279 | -0.9968 | -0.063215 | -0.9968 | -0.063215 | -0.9968 | -0.063215 | -0.010426 | -176.3712 | 0.030038 | -1.5839 | 1.5842 | -88.9135 | 0.017448
13 | 2290.8677 | 0.005 | -1.4782 | 1.4782 | -89.8062 |  | 0.0036816 | -0.058897 | 0.0036779 | -0.058838 | 0.0036779 | -0.058838 | 0.0036779 | -0.058838 | -24.5899 | -86.4231 | 0.0050823 | -1.4767 | 1.4767 | -89.8028 | -0.10002 | -0.99805 | -0.059063 | -0.99706 | -0.059004 | -0.99706 | -0.059004 | -0.99706 | -0.059004 | -0.010426 | -176.6133 | 0.030034 | -1.4782 | 1.4785 | -88.836 | 0.020033
14 | 2454.7089 | 0.005 | -1.3795 | 1.3795 | -89.7923 |  | 0.0032337 | -0.05499 | 0.0032305 | -0.054935 | 0.0032305 | -0.054935 | 0.0032305 | -0.054935 | -25.188 | -86.6346 | 0.005071 | -1.3781 | 1.3781 | -89.7892 | -0.10002 | -0.99828 | -0.055127 | -0.99728 | -0.055072 | -0.99728 | -0.055072 | -0.99728 | -0.055072 | -0.010426 | -176.8392 | 0.030032 | -1.3795 | 1.3798 | -88.7529 | 0.023002
15 | 2630.268 | 0.005 | -1.2874 | 1.2874 | -89.7775 |  | 0.0028433 | -0.05134 | 0.0028404 | -0.051289 | 0.0028404 | -0.051289 | 0.0028404 | -0.051289 | -25.7863 | -86.8301 | 0.0050612 | -1.2861 | 1.2861 | -89.7745 | -0.10002 | -0.99848 | -0.051452 | -0.99748 | -0.051401 | -0.99748 | -0.051401 | -0.99748 | -0.051401 | -0.010426 | -177.0501 | 0.030029 | -1.2874 | 1.2878 | -88.6638 | 0.02641
16 | 2818.3829 | 0.005 | -1.2015 | 1.2015 | -89.7616 |  | 0.002503 | -0.04793 | 0.0025005 | -0.047882 | 0.0025005 | -0.047882 | 0.0025005 | -0.047882 | -26.3848 | -87.0106 | 0.0050527 | -1.2003 | 1.2003 | -89.7588 | -0.10002 | -0.99865 | -0.048022 | -0.99765 | -0.047974 | -0.99765 | -0.047974 | -0.99765 | -0.047974 | -0.010426 | -177.2469 | 0.030027 | -1.2015 | 1.2019 | -88.5684 | 0.030323
17 | 3019.9517 | 0.005 | -1.1213 | 1.1213 | -89.7445 |  | 0.0022064 | -0.044744 | 0.0022042 | -0.044699 | 0.0022042 | -0.044699 | 0.0022042 | -0.044699 | -26.9835 | -87.1769 | 0.0050452 | -1.1202 | 1.1202 | -89.7419 | -0.10002 | -0.99879 | -0.04482 | -0.9978 | -0.044775 | -0.9978 | -0.044775 | -0.9978 | -0.044775 | -0.010427 | -177.4306 | 0.030025 | -1.1213 | 1.1217 | -88.4661 | 0.034815
18 | 3235.9366 | 0.005 | -1.0464 | 1.0465 | -89.7262 |  | 0.0019479 | -0.041768 | 0.001946 | -0.041726 | 0.001946 | -0.041726 | 0.001946 | -0.041726 | -27.5824 | -87.3298 | 0.0050388 | -1.0454 | 1.0454 | -89.7238 | -0.10002 | -0.99892 | -0.041831 | -0.99793 | -0.041789 | -0.99793 | -0.041789 | -0.99793 | -0.041789 | -0.010427 | -177.6021 | 0.030023 | -1.0464 | 1.0469 | -88.3566 | 0.039973
19 | 3467.3685 | 0.005 | -0.9766 | 0.97661 | -89.7067 |  | 0.0017227 | -0.038989 | 0.001721 | -0.03895 | 0.001721 | -0.03895 | 0.001721 | -0.03895 | -28.1814 | -87.47 | 0.0050331 | -0.97562 | 0.97564 | -89.7044 | -0.10002 | -0.99904 | -0.039041 | -0.99804 | -0.039002 | -0.99804 | -0.039002 | -0.99804 | -0.039002 | -0.010427 | -177.7621 | 0.030022 | -0.9766 | 0.97706 | -88.2392 | 0.045894
20 | 3715.3523 | 0.005 | -0.91142 | 0.91143 | -89.6857 |  | 0.0015265 | -0.036394 | 0.001525 | -0.036357 | 0.001525 | -0.036357 | 0.001525 | -0.036357 | -28.7805 | -87.5982 | 0.0050282 | -0.9105 | 0.91052 | -89.6836 | -0.10002 | -0.99914 | -0.036437 | -0.99814 | -0.036401 | -0.99814 | -0.036401 | -0.99814 | -0.036401 | -0.010427 | -177.9114 | 0.030021 | -0.91142 | 0.91191 | -88.1134 | 0.052692
21 | 3981.0717 | 0.005 | -0.85058 | 0.8506 | -89.6632 |  | 0.0013555 | -0.03397 | 0.0013542 | -0.033936 | 0.0013542 | -0.033936 | 0.0013542 | -0.033936 | -29.3798 | -87.715 | 0.0050239 | -0.84973 | 0.84975 | -89.6613 | -0.10002 | -0.99922 | -0.034007 | -0.99822 | -0.033973 | -0.99822 | -0.033973 | -0.99822 | -0.033973 | -0.010427 | -178.0508 | 0.03002 | -0.85058 | 0.85111 | -87.9787 | 0.060497
22 | 4265.7952 | 0.005 | -0.79381 | 0.79382 | -89.6391 |  | 0.0012065 | -0.031708 | 0.0012053 | -0.031676 | 0.0012053 | -0.031676 | 0.0012053 | -0.031676 | -29.9791 | -87.8208 | 0.0050202 | -0.79301 | 0.79303 | -89.6373 | -0.10002 | -0.9993 | -0.031738 | -0.9983 | -0.031706 | -0.9983 | -0.031706 | -0.9983 | -0.031706 | -0.010427 | -178.1809 | 0.030019 | -0.79381 | 0.79437 | -87.8343 | 0.069458
23 | 4570.8819 | 0.005 | -0.74082 | 0.74084 | -89.6133 |  | 0.0010768 | -0.029595 | 0.0010757 | -0.029565 | 0.0010757 | -0.029565 | 0.0010757 | -0.029565 | -30.5786 | -87.9163 | 0.0050169 | -0.74008 | 0.7401 | -89.6116 | -0.10002 | -0.99936 | -0.02962 | -0.99836 | -0.029591 | -0.99836 | -0.029591 | -0.99836 | -0.029591 | -0.010427 | -178.3023 | 0.030018 | -0.74082 | 0.74143 | -87.6797 | 0.079745
24 | 4897.7882 | 0.005 | -0.69137 | 0.69139 | -89.5856 |  | 0.00096371 | -0.027623 | 0.00096275 | -0.027595 | 0.00096275 | -0.027595 | 0.00096275 | -0.027595 | -31.1781 | -88.0019 | 0.0050141 | -0.69068 | 0.6907 | -89.5841 | -0.10002 | -0.99942 | -0.027644 | -0.99842 | -0.027616 | -0.99842 | -0.027616 | -0.99842 | -0.027616 | -0.010427 | -178.4156 | 0.030017 | -0.69137 | 0.69202 | -87.514 | 0.091555
25 | 5248.0746 | 0.005 | -0.64522 | 0.64524 | -89.556 |  | 0.00086522 | -0.025782 | 0.00086436 | -0.025756 | 0.00086436 | -0.025756 | 0.00086436 | -0.025756 | -31.7776 | -88.0779 | 0.0050116 | -0.64458 | 0.6446 | -89.5545 | -0.10002 | -0.99947 | -0.0258 | -0.99847 | -0.025774 | -0.99847 | -0.025774 | -0.99847 | -0.025774 | -0.010427 | -178.5213 | 0.030017 | -0.64522 | 0.64592 | -87.3364 | 0.10511
26 | 5623.4133 | 0.005 | -0.60216 | 0.60218 | -89.5243 |  | 0.00077942 | -0.024063 | 0.00077864 | -0.024039 | 0.00077864 | -0.024039 | 0.00077864 | -0.024039 | -32.3773 | -88.1448 | 0.0050095 | -0.60155 | 0.60158 | -89.5229 | -0.10002 | -0.99951 | -0.024078 | -0.99851 | -0.024054 | -0.99851 | -0.024054 | -0.99851 | -0.024054 | -0.010427 | -178.62 | 0.030016 | -0.60216 | 0.6029 | -87.1463 | 0.12068
27 | 6025.5959 | 0.005 | -0.56196 | 0.56199 | -89.4902 |  | 0.00070469 | -0.022458 | 0.00070398 | -0.022436 | 0.00070398 | -0.022436 | 0.00070398 | -0.022436 | -32.9769 | -88.2028 | 0.0050076 | -0.5614 | 0.56142 | -89.4889 | -0.10002 | -0.99955 | -0.022471 | -0.99855 | -0.022449 | -0.99855 | -0.022449 | -0.99855 | -0.022449 | -0.010427 | -178.7121 | 0.030016 | -0.56196 | 0.56276 | -86.9426 | 0.13854
28 | 6456.5423 | 0.005 | -0.52445 | 0.52448 | -89.4538 |  | 0.00063958 | -0.02096 | 0.00063894 | -0.020939 | 0.00063894 | -0.020939 | 0.00063894 | -0.020939 | -33.5766 | -88.2522 | 0.005006 | -0.52393 | 0.52395 | -89.4526 | -0.10002 | -0.99958 | -0.020972 | -0.99858 | -0.020951 | -0.99858 | -0.020951 | -0.99858 | -0.020951 | -0.010427 | -178.7981 | 0.030015 | -0.52445 | 0.52531 | -86.7244 | 0.15905
29 | 6918.3097 | 0.005 | -0.48944 | 0.48947 | -89.4147 |  | 0.00058287 | -0.019562 | 0.00058229 | -0.019543 | 0.00058229 | -0.019543 | 0.00058229 | -0.019543 | -34.1764 | -88.2933 | 0.0050046 | -0.48895 | 0.48898 | -89.4136 | -0.10002 | -0.99961 | -0.019572 | -0.99861 | -0.019552 | -0.99861 | -0.019552 | -0.99861 | -0.019552 | -0.010427 | -178.8783 | 0.030015 | -0.48944 | 0.49036 | -86.4908 | 0.1826
30 | 7413.1024 | 0.005 | -0.45677 | 0.4568 | -89.3728 |  | 0.00053347 | -0.018258 | 0.00053294 | -0.018239 | 0.00053294 | -0.018239 | 0.00053294 | -0.018239 | -34.7762 | -88.3263 | 0.0050033 | -0.45632 | 0.45634 | -89.3718 | -0.10002 | -0.99963 | -0.018266 | -0.99863 | -0.018247 | -0.99863 | -0.018247 | -0.99863 | -0.018247 | -0.010427 | -178.9532 | 0.030015 | -0.45677 | 0.45776 | -86.2405 | 0.20962
31 | 7943.2823 | 0.005 | -0.42628 | 0.42631 | -89.328 |  | 0.00049045 | -0.01704 | 0.00048996 | -0.017022 | 0.00048996 | -0.017022 | 0.00048996 | -0.017022 | -35.376 | -88.3513 | 0.0050023 | -0.42586 | 0.42588 | -89.327 | -0.10002 | -0.99965 | -0.017047 | -0.99866 | -0.01703 | -0.99866 | -0.01703 | -0.99866 | -0.01703 | -0.010427 | -179.0231 | 0.030014 | -0.42628 | 0.42734 | -85.9725 | 0.24064
32 | 8511.3804 | 0.005 | -0.39783 | 0.39786 | -89.2799 |  | 0.00045297 | -0.015903 | 0.00045252 | -0.015887 | 0.00045252 | -0.015887 | 0.00045252 | -0.015887 | -35.9758 | -88.3684 | 0.0050013 | -0.39743 | 0.39746 | -89.279 | -0.10002 | -0.99967 | -0.015909 | -0.99867 | -0.015893 | -0.99867 | -0.015893 | -0.99867 | -0.015893 | -0.010427 | -179.0883 | 0.030014 | -0.39783 | 0.39896 | -85.6855 | 0.27624
33 | 9120.1084 | 0.005 | -0.37127 | 0.3713 | -89.2284 |  | 0.00042032 | -0.014842 | 0.0004199 | -0.014827 | 0.0004199 | -0.014827 | 0.0004199 | -0.014827 | -36.5756 | -88.3778 | 0.0050005 | -0.3709 | 0.37093 | -89.2276 | -0.10002 | -0.99969 | -0.014847 | -0.99869 | -0.014832 | -0.99869 | -0.014832 | -0.99869 | -0.014832 | -0.010427 | -179.1491 | 0.030014 | -0.37127 | 0.37248 | -85.3782 | 0.3171
34 | 9772.3722 | 0.005 | -0.34648 | 0.34652 | -89.1732 |  | 0.00039189 | -0.013851 | 0.0003915 | -0.013837 | 0.0003915 | -0.013837 | 0.0003915 | -0.013837 | -37.1755 | -88.3794 | 0.0049998 | -0.34614 | 0.34617 | -89.1724 | -0.10002 | -0.9997 | -0.013856 | -0.9987 | -0.013842 | -0.9987 | -0.013842 | -0.9987 | -0.013842 | -0.010427 | -179.2059 | 0.030014 | -0.34648 | 0.34778 | -85.0492 | 0.364
35 | 10471.2855 | 0.005 | -0.32335 | 0.32339 | -89.1141 |  | 0.00036712 | -0.012927 | 0.00036676 | -0.012914 | 0.00036676 | -0.012914 | 0.00036676 | -0.012914 | -37.7754 | -88.3732 | 0.0049992 | -0.32303 | 0.32307 | -89.1134 | -0.10002 | -0.99972 | -0.012931 | -0.99872 | -0.012918 | -0.99872 | -0.012918 | -0.99872 | -0.012918 | -0.010427 | -179.2589 | 0.030014 | -0.32335 | 0.32474 | -84.697 | 0.41781
36 | 11220.1845 | 0.005 | -0.30177 | 0.30181 | -89.0507 |  | 0.00034555 | -0.012064 | 0.00034521 | -0.012052 | 0.00034521 | -0.012052 | 0.00034521 | -0.012052 | -38.3752 | -88.3593 | 0.0049986 | -0.30146 | 0.30151 | -89.0501 | -0.10002 | -0.99973 | -0.012068 | -0.99873 | -0.012056 | -0.99873 | -0.012056 | -0.99873 | -0.012056 | -0.010427 | -179.3084 | 0.030013 | -0.30177 | 0.30326 | -84.3201 | 0.47956
37 | 12022.6443 | 0.005 | -0.28162 | 0.28166 | -88.9829 |  | 0.00032676 | -0.011259 | 0.00032644 | -0.011248 | 0.00032644 | -0.011248 | 0.00032644 | -0.011248 | -38.9751 | -88.3376 | 0.0049982 | -0.28134 | 0.28138 | -88.9822 | -0.10002 | -0.99974 | -0.011262 | -0.99874 | -0.011251 | -0.99874 | -0.011251 | -0.99874 | -0.011251 | -0.010427 | -179.3546 | 0.030013 | -0.28162 | 0.28322 | -83.9167 | 0.55041
38 | 12882.4955 | 0.005 | -0.26282 | 0.26287 | -88.9101 |  | 0.0003104 | -0.010507 | 0.00031009 | -0.010497 | 0.00031009 | -0.010497 | 0.00031009 | -0.010497 | -39.575 | -88.3079 | 0.0049978 | -0.26256 | 0.2626 | -88.9095 | -0.10002 | -0.99974 | -0.01051 | -0.99875 | -0.0105 | -0.99875 | -0.0105 | -0.99875 | -0.0105 | -0.010427 | -179.3977 | 0.030013 | -0.26282 | 0.26453 | -83.4852 | 0.6317
39 | 13803.8426 | 0.005 | -0.24527 | 0.24532 | -88.8322 |  | 0.00029615 | -0.009806 | 0.00029585 | -0.0097962 | 0.00029585 | -0.0097962 | 0.00029585 | -0.0097962 | -40.1749 | -88.2702 | 0.0049974 | -0.24503 | 0.24508 | -88.8316 | -0.10002 | -0.99975 | -0.0098086 | -0.99875 | -0.0097988 | -0.99875 | -0.0097988 | -0.99875 | -0.0097988 | -0.010427 | -179.4379 | 0.030013 | -0.24527 | 0.2471 | -83.0235 | 0.72494
40 | 14791.0839 | 0.005 | -0.22889 | 0.22895 | -88.7486 |  | 0.00028373 | -0.0091513 | 0.00028345 | -0.0091422 | 0.00028345 | -0.0091422 | 0.00028345 | -0.0091422 | -40.7748 | -88.2242 | 0.0049971 | -0.22866 | 0.22872 | -88.7481 | -0.10002 | -0.99976 | -0.0091537 | -0.99876 | -0.0091446 | -0.99876 | -0.0091446 | -0.99876 | -0.0091446 | -0.010427 | -179.4754 | 0.030013 | -0.22889 | 0.23085 | -82.5299 | 0.8319
41 | 15848.9319 | 0.005 | -0.21361 | 0.21367 | -88.6591 |  | 0.00027292 | -0.0085403 | 0.00027264 | -0.0085318 | 0.00027264 | -0.0085318 | 0.00027264 | -0.0085318 | -41.3747 | -88.1697 | 0.0049968 | -0.2134 | 0.21345 | -88.6586 | -0.10002 | -0.99976 | -0.0085425 | -0.99876 | -0.008534 | -0.99876 | -0.008534 | -0.99876 | -0.008534 | -0.010427 | -179.5104 | 0.030013 | -0.21361 | 0.21571 | -82.0021 | 0.95455
42 | 16982.4365 | 0.005 | -0.19935 | 0.19941 | -88.5632 |  | 0.0002635 | -0.0079701 | 0.00026324 | -0.0079621 | 0.00026324 | -0.0079621 | 0.00026324 | -0.0079621 | -41.9747 | -88.1064 | 0.0049966 | -0.19915 | 0.19921 | -88.5627 | -0.10002 | -0.99977 | -0.0079721 | -0.99877 | -0.0079641 | -0.99877 | -0.0079641 | -0.99877 | -0.0079641 | -0.010427 | -179.5431 | 0.030013 | -0.19935 | 0.20159 | -81.438 | 1.0952
43 | 18197.0086 | 0.005 | -0.18603 | 0.1861 | -88.4604 |  | 0.0002553 | -0.0074379 | 0.00025504 | -0.0074305 | 0.00025504 | -0.0074305 | 0.00025504 | -0.0074305 | -42.5746 | -88.0342 | 0.0049964 | -0.18585 | 0.18591 | -88.46 | -0.10002 | -0.99977 | -0.0074397 | -0.99877 | -0.0074323 | -0.99877 | -0.0074323 | -0.99877 | -0.0074323 | -0.010427 | -179.5736 | 0.030013 | -0.18603 | 0.18844 | -80.8354 | 1.2564
44 | 19498.446 | 0.005 | -0.17361 | 0.17368 | -88.3503 |  | 0.00024815 | -0.0069412 | 0.0002479 | -0.0069343 | 0.0002479 | -0.0069343 | 0.0002479 | -0.0069343 | -43.1745 | -87.9525 | 0.0049962 | -0.17343 | 0.17351 | -88.3499 | -0.10002 | -0.99978 | -0.0069428 | -0.99878 | -0.0069359 | -0.99878 | -0.0069359 | -0.99878 | -0.0069359 | -0.010427 | -179.6021 | 0.030013 | -0.17361 | 0.17618 | -80.1918 | 1.4412
45 | 20892.9613 | 0.005 | -0.16201 | 0.16209 | -88.2323 |  | 0.00024193 | -0.0064776 | 0.00024169 | -0.0064711 | 0.00024169 | -0.0064711 | 0.00024169 | -0.0064711 | -43.7743 | -87.8611 | 0.004996 | -0.16185 | 0.16193 | -88.2319 | -0.10002 | -0.99978 | -0.0064791 | -0.99878 | -0.0064726 | -0.99878 | -0.0064726 | -0.99878 | -0.0064726 | -0.010427 | -179.6287 | 0.030013 | -0.16201 | 0.16477 | -79.5049 | 1.653
46 | 22387.2114 | 0.005 | -0.15119 | 0.15127 | -88.1059 |  | 0.00023651 | -0.0060449 | 0.00023627 | -0.0060389 | 0.00023627 | -0.0060389 | 0.00023627 | -0.0060389 | -44.3742 | -87.7594 | 0.0049959 | -0.15104 | 0.15112 | -88.1055 | -0.10002 | -0.99978 | -0.0060463 | -0.99878 | -0.0060403 | -0.99878 | -0.0060403 | -0.99878 | -0.0060403 | -0.010427 | -179.6535 | 0.030013 | -0.15119 | 0.15414 | -78.7721 | 1.8956
47 | 23988.3292 | 0.005 | -0.14109 | 0.14118 | -87.9704 |  | 0.00023179 | -0.0056411 | 0.00023156 | -0.0056354 | 0.00023156 | -0.0056354 | 0.00023156 | -0.0056354 | -44.9741 | -87.6471 | 0.0049958 | -0.14095 | 0.14104 | -87.97 | -0.10002 | -0.99978 | -0.0056424 | -0.99878 | -0.0056367 | -0.99878 | -0.0056367 | -0.99878 | -0.0056367 | -0.010427 | -179.6767 | 0.030013 | -0.14109 | 0.14424 | -77.9909 | 2.1733
48 | 25703.9578 | 0.005 | -0.13166 | 0.13176 | -87.8252 |  | 0.00022768 | -0.0052642 | 0.00022745 | -0.0052589 | 0.00022745 | -0.0052589 | 0.00022745 | -0.0052589 | -45.574 | -87.5235 | 0.0049957 | -0.13153 | 0.13162 | -87.8249 | -0.10002 | -0.99979 | -0.0052653 | -0.99879 | -0.0052601 | -0.99879 | -0.0052601 | -0.99879 | -0.0052601 | -0.010427 | -179.6983 | 0.030013 | -0.13166 | 0.13504 | -77.1586 | 2.4914
49 | 27542.287 | 0.005 | -0.12286 | 0.12296 | -87.6696 |  | 0.0002241 | -0.0049124 | 0.00022387 | -0.0049075 | 0.00022387 | -0.0049075 | 0.00022387 | -0.0049075 | -46.1738 | -87.388 | 0.0049956 | -0.12274 | 0.12284 | -87.6693 | -0.10002 | -0.99979 | -0.0049135 | -0.99879 | -0.0049085 | -0.99879 | -0.0049085 | -0.99879 | -0.0049085 | -0.010427 | -179.7184 | 0.030013 | -0.12286 | 0.12647 | -76.2726 | 2.8552
50 | 29512.0923 | 0.005 | -0.11465 | 0.11476 | -87.5028 |  | 0.00022098 | -0.004584 | 0.00022076 | -0.0045795 | 0.00022076 | -0.0045795 | 0.00022076 | -0.0045795 | -46.7736 | -87.2401 | 0.0049955 | -0.11453 | 0.11464 | -87.5026 | -0.10002 | -0.99979 | -0.004585 | -0.99879 | -0.0045804 | -0.99879 | -0.0045804 | -0.99879 | -0.0045804 | -0.010427 | -179.7372 | 0.030013 | -0.11465 | 0.11851 | -75.3304 | 3.2714
51 | 31622.7766 | 0.005 | -0.10698 | 0.1071 | -87.3242 |  | 0.00021826 | -0.0042776 | 0.00021804 | -0.0042733 | 0.00021804 | -0.0042733 | 0.00021804 | -0.0042733 | -47.3734 | -87.079 | 0.0049955 | -0.10688 | 0.10699 | -87.3239 | -0.10002 | -0.99979 | -0.0042785 | -0.99879 | -0.0042742 | -0.99879 | -0.0042742 | -0.99879 | -0.0042742 | -0.010427 | -179.7548 | 0.030013 | -0.10698 | 0.11111 | -74.3294 | 3.7472
52 | 33884.4156 | 0.005 | -0.09983 | 0.099955 | -87.1327 |  | 0.0002159 | -0.0039915 | 0.00021568 | -0.0039875 | 0.00021568 | -0.0039875 | 0.00021568 | -0.0039875 | -47.9732 | -86.904 | 0.0049954 | -0.09973 | 0.099855 | -87.1325 | -0.10002 | -0.99979 | -0.0039924 | -0.99879 | -0.0039884 | -0.99879 | -0.0039884 | -0.99879 | -0.0039884 | -0.010427 | -179.7712 | 0.030013 | -0.09983 | 0.10424 | -73.2672 | 4.2907
53 | 36307.8055 | 0.005 | -0.093152 | 0.093286 | -86.9275 |  | 0.00021384 | -0.0037245 | 0.00021362 | -0.0037208 | 0.00021362 | -0.0037208 | 0.00021362 | -0.0037208 | -48.573 | -86.7141 | 0.0049953 | -0.093059 | 0.093192 | -86.9273 | -0.10002 | -0.99979 | -0.0037253 | -0.99879 | -0.0037216 | -0.99879 | -0.0037216 | -0.99879 | -0.0037216 | -0.010427 | -179.7865 | 0.030013 | -0.093152 | 0.097867 | -72.1415 | 4.9112
54 | 38904.5145 | 0.005 | -0.086918 | 0.087062 | -86.7077 |  | 0.00021204 | -0.0034753 | 0.00021183 | -0.0034718 | 0.00021183 | -0.0034718 | 0.00021183 | -0.0034718 | -49.1727 | -86.5085 | 0.0049953 | -0.086832 | 0.086975 | -86.7075 | -0.10002 | -0.99979 | -0.003476 | -0.99879 | -0.0034726 | -0.99879 | -0.0034726 | -0.99879 | -0.0034726 | -0.010427 | -179.8008 | 0.030013 | -0.086918 | 0.091954 | -70.9504 | 5.619
55 | 41686.9383 | 0.005 | -0.0811 | 0.081254 | -86.4721 |  | 0.00021048 | -0.0032427 | 0.00021027 | -0.0032394 | 0.00021027 | -0.0032394 | 0.00021027 | -0.0032394 | -49.7724 | -86.2862 | 0.0049953 | -0.081019 | 0.081173 | -86.4719 | -0.10002 | -0.99979 | -0.0032433 | -0.99879 | -0.0032401 | -0.99879 | -0.0032401 | -0.99879 | -0.0032401 | -0.010427 | -179.8141 | 0.030013 | -0.0811 | 0.086475 | -69.6921 | 6.4258
56 | 44668.3592 | 0.005 | -0.075669 | 0.075834 | -86.2195 |  | 0.00020912 | -0.0030255 | 0.00020891 | -0.0030225 | 0.00020891 | -0.0030225 | 0.00020891 | -0.0030225 | -50.372 | -86.0462 | 0.0049952 | -0.075593 | 0.075758 | -86.2194 | -0.10002 | -0.9998 | -0.0030261 | -0.9988 | -0.0030231 | -0.9988 | -0.0030231 | -0.9988 | -0.0030231 | -0.010427 | -179.8266 | 0.030013 | -0.075669 | 0.081404 | -68.3652 | 7.3445
57 | 47863.0092 | 0.005 | -0.070599 | 0.070776 | -85.9489 |  | 0.00020793 | -0.0028228 | 0.00020772 | -0.00282 | 0.00020772 | -0.00282 | 0.00020772 | -0.00282 | -50.9716 | -85.7872 | 0.0049952 | -0.070528 | 0.070705 | -85.9488 | -0.10002 | -0.9998 | -0.0028234 | -0.9988 | -0.0028206 | -0.9988 | -0.0028206 | -0.9988 | -0.0028206 | -0.010427 | -179.8382 | 0.030013 | -0.070599 | 0.076714 | -66.969 | 8.3895
58 | 51286.1384 | 0.005 | -0.065866 | 0.066055 | -85.6589 |  | 0.0002069 | -0.0026336 | 0.00020669 | -0.0026309 | 0.00020669 | -0.0026309 | 0.00020669 | -0.0026309 | -51.5711 | -85.508 | 0.0049952 | -0.0658 | 0.065989 | -85.6588 | -0.10002 | -0.9998 | -0.0026341 | -0.9988 | -0.0026315 | -0.9988 | -0.0026315 | -0.9988 | -0.0026315 | -0.010427 | -179.849 | 0.030013 | -0.065866 | 0.072381 | -65.5031 | 9.5768
59 | 54954.0874 | 0.005 | -0.061447 | 0.061651 | -85.3481 |  | 0.000206 | -0.0024569 | 0.00020579 | -0.0024544 | 0.00020579 | -0.0024544 | 0.00020579 | -0.0024544 | -52.1705 | -85.2073 | 0.0049951 | -0.061386 | 0.061589 | -85.3479 | -0.10002 | -0.9998 | -0.0024574 | -0.9988 | -0.0024549 | -0.9988 | -0.0024549 | -0.9988 | -0.0024549 | -0.010427 | -179.8592 | 0.030013 | -0.061447 | 0.068385 | -63.9679 | 10.924
60 | 58884.3655 | 0.005 | -0.057322 | 0.05754 | -85.0149 |  | 0.00020521 | -0.002292 | 0.00020501 | -0.0022897 | 0.00020501 | -0.0022897 | 0.00020501 | -0.0022897 | -52.7699 | -84.8836 | 0.0049951 | -0.057265 | 0.057482 | -85.0148 | -0.10002 | -0.9998 | -0.0022924 | -0.9988 | -0.0022901 | -0.9988 | -0.0022901 | -0.9988 | -0.0022901 | -0.010427 | -179.8686 | 0.030013 | -0.057322 | 0.064704 | -62.3645 | 12.4504
61 | 63095.7344 | 0.005 | -0.053471 | 0.053704 | -84.6578 |  | 0.00020453 | -0.002138 | 0.00020433 | -0.0021358 | 0.00020433 | -0.0021358 | 0.00020433 | -0.0021358 | -53.3691 | -84.5353 | 0.0049951 | -0.053417 | 0.05365 | -84.6577 | -0.10002 | -0.9998 | -0.0021384 | -0.9988 | -0.0021363 | -0.9988 | -0.0021363 | -0.9988 | -0.0021363 | -0.010427 | -179.8775 | 0.030013 | -0.053471 | 0.061318 | -60.6949 | 14.1774
62 | 67608.2975 | 0.005 | -0.049874 | 0.050124 | -84.2751 |  | 0.00020394 | -0.0019942 | 0.00020373 | -0.0019922 | 0.00020373 | -0.0019922 | 0.00020373 | -0.0019922 | -53.9683 | -84.1608 | 0.0049951 | -0.049824 | 0.050074 | -84.275 | -0.10002 | -0.9998 | -0.0019946 | -0.9988 | -0.0019926 | -0.9988 | -0.0019926 | -0.9988 | -0.0019926 | -0.010427 | -179.8857 | 0.030013 | -0.049874 | 0.058208 | -58.962 | 16.1277
63 | 72443.596 | 0.005 | -0.046516 | 0.046784 | -83.8648 |  | 0.00020342 | -0.0018599 | 0.00020322 | -0.001858 | 0.00020322 | -0.001858 | 0.00020322 | -0.001858 | -54.5673 | -83.7583 | 0.0049951 | -0.046469 | 0.046737 | -83.8647 | -0.10002 | -0.9998 | -0.0018603 | -0.9988 | -0.0018584 | -0.9988 | -0.0018584 | -0.9988 | -0.0018584 | -0.010427 | -179.8934 | 0.030013 | -0.046516 | 0.055358 | -57.1695 | 18.3265
64 | 77624.7117 | 0.005 | -0.04338 | 0.043667 | -83.425 |  | 0.00020297 | -0.0017345 | 0.00020277 | -0.0017328 | 0.00020277 | -0.0017328 | 0.00020277 | -0.0017328 | -55.1662 | -83.3256 | 0.0049951 | -0.043336 | 0.043623 | -83.4249 | -0.10002 | -0.9998 | -0.0017348 | -0.9988 | -0.0017331 | -0.9988 | -0.0017331 | -0.9988 | -0.0017331 | -0.010427 | -179.9006 | 0.030013 | -0.04338 | 0.05275 | -55.3224 | 20.8005
65 | 83176.3771 | 0.005 | -0.040451 | 0.040759 | -82.9536 |  | 0.00020258 | -0.0016174 | 0.00020237 | -0.0016158 | 0.00020237 | -0.0016158 | 0.00020237 | -0.0016158 | -55.7649 | -82.8609 | 0.0049951 | -0.04041 | 0.040718 | -82.9535 | -0.10002 | -0.9998 | -0.0016177 | -0.9988 | -0.0016161 | -0.9988 | -0.0016161 | -0.9988 | -0.0016161 | -0.010427 | -179.9073 | 0.030013 | -0.040451 | 0.050369 | -53.4263 | 23.5784
66 | 89125.0938 | 0.005 | -0.037715 | 0.038045 | -82.4481 |  | 0.00020223 | -0.001508 | 0.00020203 | -0.0015065 | 0.00020203 | -0.0015065 | 0.00020203 | -0.0015065 | -56.3634 | -82.3617 | 0.0049951 | -0.037677 | 0.038007 | -82.448 | -0.10002 | -0.9998 | -0.0015083 | -0.9988 | -0.0015068 | -0.9988 | -0.0015068 | -0.9988 | -0.0015068 | -0.010427 | -179.9136 | 0.030013 | -0.037715 | 0.048199 | -51.488 | 26.6907
67 | 95499.2586 | 0.005 | -0.035159 | 0.035512 | -81.9061 |  | 0.00020194 | -0.0014058 | 0.00020173 | -0.0014044 | 0.00020173 | -0.0014044 | 0.00020173 | -0.0014044 | -56.9616 | -81.8255 | 0.004995 | -0.035123 | 0.035477 | -81.906 | -0.10002 | -0.9998 | -0.0014061 | -0.9988 | -0.0014047 | -0.9988 | -0.0014047 | -0.9988 | -0.0014047 | -0.010427 | -179.9194 | 0.030013 | -0.035159 | 0.046226 | -49.5149 | 30.1698
68 | 102329.2992 | 0.005 | -0.03277 | 0.03315 | -81.3249 |  | 0.00020168 | -0.0013103 | 0.00020148 | -0.001309 | 0.00020148 | -0.001309 | 0.00020148 | -0.001309 | -57.5596 | -81.2498 | 0.004995 | -0.032738 | 0.033117 | -81.3248 | -0.10002 | -0.9998 | -0.0013106 | -0.9988 | -0.0013092 | -0.9988 | -0.0013092 | -0.9988 | -0.0013092 | -0.010427 | -179.9249 | 0.030013 | -0.03277 | 0.044437 | -47.5153 | 34.0497
69 | 109647.8196 | 0.005 | -0.030539 | 0.030945 | -80.7017 |  | 0.00020145 | -0.0012211 | 0.00020125 | -0.0012198 | 0.00020125 | -0.0012198 | 0.00020125 | -0.0012198 | -58.1573 | -80.6317 | 0.004995 | -0.030508 | 0.030914 | -80.7016 | -0.10002 | -0.9998 | -0.0012213 | -0.9988 | -0.0012201 | -0.9988 | -0.0012201 | -0.9988 | -0.0012201 | -0.010427 | -179.93 | 0.030013 | -0.030539 | 0.042818 | -45.4979 | 38.3659
70 | 117489.7555 | 0.005 | -0.028453 | 0.028889 | -80.0332 |  | 0.00020125 | -0.0011377 | 0.00020105 | -0.0011365 | 0.00020105 | -0.0011365 | 0.00020105 | -0.0011365 | -58.7547 | -79.968 | 0.004995 | -0.028424 | 0.02886 | -80.0331 | -0.10002 | -0.9998 | -0.0011379 | -0.9988 | -0.0011367 | -0.9988 | -0.0011367 | -0.9988 | -0.0011367 | -0.010427 | -179.9348 | 0.030013 | -0.028453 | 0.041356 | -43.4718 | 43.1557
71 | 125892.5412 | 0.005 | -0.026503 | 0.02697 | -79.3161 |  | 0.00020108 | -0.0010597 | 0.00020088 | -0.0010586 | 0.00020088 | -0.0010586 | 0.00020088 | -0.0010586 | -59.3516 | -79.2554 | 0.004995 | -0.026476 | 0.026943 | -79.3161 | -0.10002 | -0.9998 | -0.0010599 | -0.9988 | -0.0010588 | -0.9988 | -0.0010588 | -0.9988 | -0.0010588 | -0.010427 | -179.9393 | 0.030013 | -0.026503 | 0.040039 | -41.4462 | 48.4575
72 | 134896.2883 | 0.005 | -0.024679 | 0.02518 | -78.5468 |  | 0.00020093 | -0.00098676 | 0.00020073 | -0.00098578 | 0.00020073 | -0.00098578 | 0.00020073 | -0.00098578 | -59.948 | -78.4903 | 0.004995 | -0.024654 | 0.025155 | -78.5468 | -0.10002 | -0.9998 | -0.00098696 | -0.9988 | -0.00098598 | -0.9988 | -0.00098598 | -0.9988 | -0.00098598 | -0.010427 | -179.9434 | 0.030013 | -0.024679 | 0.038856 | -39.4301 | 54.3111
73 | 144543.9771 | 0.005 | -0.022973 | 0.023511 | -77.7214 |  | 0.0002008 | -0.00091856 | 0.0002006 | -0.00091764 | 0.0002006 | -0.00091764 | 0.0002006 | -0.00091764 | -60.5438 | -77.6687 | 0.004995 | -0.02295 | 0.023488 | -77.7213 | -0.10002 | -0.9998 | -0.00091874 | -0.9988 | -0.00091783 | -0.9988 | -0.00091783 | -0.9988 | -0.00091783 | -0.010427 | -179.9473 | 0.030013 | -0.022973 | 0.037796 | -37.4324 | 60.7576
74 | 154881.6619 | 0.005 | -0.021377 | 0.021954 | -76.8354 |  | 0.00020069 | -0.00085474 | 0.00020049 | -0.00085388 | 0.00020049 | -0.00085388 | 0.00020049 | -0.00085388 | -61.139 | -76.7864 | 0.004995 | -0.021356 | 0.021932 | -76.8354 | -0.10002 | -0.9998 | -0.00085491 | -0.9988 | -0.00085406 | -0.9988 | -0.00085406 | -0.9988 | -0.00085406 | -0.010427 | -179.951 | 0.030013 | -0.021377 | 0.036847 | -35.4612 | 67.8389
75 | 165958.6907 | 0.005 | -0.019883 | 0.020502 | -75.8844 |  | 0.00020059 | -0.000795 | 0.00020039 | -0.0007942 | 0.00020039 | -0.0007942 | 0.00020039 | -0.0007942 | -61.7333 | -75.8389 | 0.004995 | -0.019863 | 0.020481 | -75.8844 | -0.10002 | -0.9998 | -0.00079516 | -0.9988 | -0.00079436 | -0.9988 | -0.00079436 | -0.9988 | -0.00079436 | -0.010427 | -179.9544 | 0.030013 | -0.019883 | 0.036001 | -33.524 | 75.5982
76 | 177827.941 | 0.005 | -0.018484 | 0.019148 | -74.8633 |  | 0.00020051 | -0.00073905 | 0.00020031 | -0.00073832 | 0.00020031 | -0.00073832 | 0.00020031 | -0.00073832 | -62.3267 | -74.821 | 0.004995 | -0.018465 | 0.019129 | -74.8633 | -0.10002 | -0.9998 | -0.0007392 | -0.9988 | -0.00073846 | -0.9988 | -0.00073846 | -0.9988 | -0.00073846 | -0.010427 | -179.9576 | 0.030013 | -0.018484 | 0.035248 | -31.6276 | 84.0794
77 | 190546.0718 | 0.005 | -0.017173 | 0.017886 | -73.7667 |  | 0.00020043 | -0.00068664 | 0.00020023 | -0.00068595 | 0.00020023 | -0.00068595 | 0.00020023 | -0.00068595 | -62.919 | -73.7273 | 0.004995 | -0.017156 | 0.017868 | -73.7666 | -0.10002 | -0.9998 | -0.00068677 | -0.9988 | -0.00068609 | -0.9988 | -0.00068609 | -0.9988 | -0.00068609 | -0.010427 | -179.9606 | 0.030013 | -0.017173 | 0.034578 | -29.7777 | 93.3269
78 | 204173.7945 | 0.005 | -0.015944 | 0.016709 | -72.5885 |  | 0.00020037 | -0.0006375 | 0.00020017 | -0.00063686 | 0.00020017 | -0.00063686 | 0.00020017 | -0.00063686 | -63.51 | -72.552 | 0.004995 | -0.015928 | 0.016693 | -72.5885 | -0.10002 | -0.9998 | -0.00063763 | -0.9988 | -0.00063699 | -0.9988 | -0.00063699 | -0.9988 | -0.00063699 | -0.010427 | -179.9635 | 0.030013 | -0.015944 | 0.033985 | -27.979 | 103.3858
79 | 218776.1624 | 0.005 | -0.014791 | 0.015613 | -71.3225 |  | 0.00020031 | -0.0005914 | 0.00020011 | -0.00059081 | 0.00020011 | -0.00059081 | 0.00020011 | -0.00059081 | -64.0994 | -71.2886 | 0.004995 | -0.014776 | 0.015598 | -71.3224 | -0.10002 | -0.9998 | -0.00059152 | -0.9988 | -0.00059093 | -0.9988 | -0.00059093 | -0.9988 | -0.00059093 | -0.010427 | -179.9661 | 0.030013 | -0.014791 | 0.033459 | -26.2353 | 114.301
80 | 234422.8815 | 0.005 | -0.013709 | 0.014592 | -69.9615 |  | 0.00020026 | -0.00054813 | 0.00020006 | -0.00054758 | 0.00020006 | -0.00054758 | 0.00020006 | -0.00054758 | -64.6869 | -69.9301 | 0.004995 | -0.013695 | 0.014577 | -69.9614 | -0.10002 | -0.9998 | -0.00054824 | -0.9988 | -0.00054769 | -0.9988 | -0.00054769 | -0.9988 | -0.00054769 | -0.010427 | -179.9686 | 0.030013 | -0.013709 | 0.032995 | -24.5493 | 126.117
81 | 251188.6432 | 0.005 | -0.012692 | 0.013641 | -68.4979 |  | 0.00020022 | -0.00050747 | 0.00020002 | -0.00050696 | 0.00020002 | -0.00050696 | 0.00020002 | -0.00050696 | -65.2721 | -68.4688 | 0.004995 | -0.012679 | 0.013628 | -68.4979 | -0.10002 | -0.9998 | -0.00050757 | -0.9988 | -0.00050707 | -0.9988 | -0.00050707 | -0.9988 | -0.00050707 | -0.010427 | -179.9709 | 0.030013 | -0.012692 | 0.032586 | -22.9228 | 138.8769
82 | 269153.4804 | 0.005 | -0.011736 | 0.012756 | -66.9234 |  | 0.00020018 | -0.00046924 | 0.00019998 | -0.00046877 | 0.00019998 | -0.00046877 | 0.00019998 | -0.00046877 | -65.8547 | -66.8965 | 0.004995 | -0.011724 | 0.012744 | -66.9234 | -0.10002 | -0.9998 | -0.00046933 | -0.9988 | -0.00046886 | -0.9988 | -0.00046886 | -0.9988 | -0.00046886 | -0.010427 | -179.9731 | 0.030013 | -0.011736 | 0.032225 | -21.3567 | 152.6218
83 | 288403.1503 | 0.005 | -0.010835 | 0.011933 | -65.2291 |  | 0.00020015 | -0.00043324 | 0.00019995 | -0.00043281 | 0.00019995 | -0.00043281 | 0.00019995 | -0.00043281 | -66.4339 | -65.2042 | 0.004995 | -0.010825 | 0.011921 | -65.229 | -0.10002 | -0.9998 | -0.00043333 | -0.9988 | -0.0004329 | -0.9988 | -0.0004329 | -0.9988 | -0.0004329 | -0.010427 | -179.9752 | 0.030013 | -0.010835 | 0.031909 | -19.8511 | 167.3885
84 | 309029.5433 | 0.005 | -0.0099869 | 0.011169 | -63.405 |  | 0.00020012 | -0.00039932 | 0.00019992 | -0.00039892 | 0.00019992 | -0.00039892 | 0.00019992 | -0.00039892 | -67.0092 | -63.3821 | 0.004995 | -0.0099769 | 0.011157 | -63.4049 | -0.10002 | -0.9998 | -0.0003994 | -0.9988 | -0.000399 | -0.9988 | -0.000399 | -0.9988 | -0.000399 | -0.010427 | -179.9771 | 0.030013 | -0.0099869 | 0.031631 | -18.4053 | 183.208
85 | 331131.1215 | 0.005 | -0.0091861 | 0.010459 | -61.4405 |  | 0.00020009 | -0.0003673 | 0.00019989 | -0.00036693 | 0.00019989 | -0.00036693 | 0.00019989 | -0.00036693 | -67.5797 | -61.4195 | 0.004995 | -0.0091769 | 0.010448 | -61.4405 | -0.10002 | -0.9998 | -0.00036737 | -0.9988 | -0.000367 | -0.9988 | -0.000367 | -0.9988 | -0.000367 | -0.010427 | -179.9789 | 0.030013 | -0.0091861 | 0.031387 | -17.0181 | 200.1026
86 | 354813.3892 | 0.005 | -0.0084291 | 0.0098005 | -59.3244 |  | 0.00020007 | -0.00033703 | 0.00019987 | -0.00033669 | 0.00019987 | -0.00033669 | 0.00019987 | -0.00033669 | -68.1442 | -59.3051 | 0.004995 | -0.0084207 | 0.0097907 | -59.3244 | -0.10002 | -0.9998 | -0.0003371 | -0.9988 | -0.00033676 | -0.9988 | -0.00033676 | -0.9988 | -0.00033676 | -0.010427 | -179.9807 | 0.030013 | -0.0084291 | 0.031174 | -15.6876 | 218.0819
87 | 380189.3963 | 0.005 | -0.0077124 | 0.0091914 | -57.0444 |  | 0.00020006 | -0.00030837 | 0.00019986 | -0.00030806 | 0.00019986 | -0.00030806 | 0.00019986 | -0.00030806 | -68.7016 | -57.0268 | 0.004995 | -0.0077047 | 0.0091822 | -57.0444 | -0.10002 | -0.9998 | -0.00030843 | -0.9988 | -0.00030813 | -0.9988 | -0.00030813 | -0.9988 | -0.00030813 | -0.010427 | -179.9823 | 0.030013 | -0.0077124 | 0.030988 | -14.4117 | 237.138
88 | 407380.2778 | 0.005 | -0.0070325 | 0.0086288 | -54.5877 |  | 0.00020004 | -0.00028119 | 0.00019984 | -0.00028091 | 0.00019984 | -0.00028091 | 0.00019984 | -0.00028091 | -69.2502 | -54.5716 | 0.004995 | -0.0070255 | 0.0086202 | -54.5877 | -0.10002 | -0.9998 | -0.00028124 | -0.9988 | -0.00028096 | -0.9988 | -0.00028096 | -0.9988 | -0.00028096 | -0.010427 | -179.9839 | 0.030013 | -0.0070325 | 0.030825 | -13.1876 | 257.2392
89 | 436515.8322 | 0.005 | -0.0063862 | 0.0081107 | -51.941 |  | 0.00020003 | -0.00025534 | 0.00019983 | -0.00025509 | 0.00019983 | -0.00025509 | 0.00019983 | -0.00025509 | -69.7881 | -51.9264 | 0.004995 | -0.0063798 | 0.0081026 | -51.941 | -0.10002 | -0.9998 | -0.0002554 | -0.9988 | -0.00025514 | -0.9988 | -0.00025514 | -0.9988 | -0.00025514 | -0.010427 | -179.9854 | 0.030013 | -0.0063862 | 0.030684 | -12.0124 | 278.3216
90 | 467735.1413 | 0.005 | -0.0057703 | 0.0076352 | -49.0908 |  | 0.00020001 | -0.00023072 | 0.00019981 | -0.00023049 | 0.00019981 | -0.00023049 | 0.00019981 | -0.00023049 | -70.3128 | -49.0776 | 0.004995 | -0.0057645 | 0.0076276 | -49.0908 | -0.10002 | -0.9998 | -0.00023077 | -0.9988 | -0.00023053 | -0.9988 | -0.00023053 | -0.9988 | -0.00023053 | -0.010427 | -179.9868 | 0.030013 | -0.0057703 | 0.030562 | -10.8831 | 300.2799
91 | 501187.2336 | 0.005 | -0.005182 | 0.0072009 | -46.0239 |  | 0.0002 | -0.0002072 | 0.0001998 | -0.00020699 | 0.0001998 | -0.00020699 | 0.0001998 | -0.00020699 | -70.8215 | -46.0121 | 0.004995 | -0.0051768 | 0.0071937 | -46.0239 | -0.10002 | -0.9998 | -0.00020724 | -0.9988 | -0.00020703 | -0.9988 | -0.00020703 | -0.9988 | -0.00020703 | -0.010427 | -179.9881 | 0.030013 | -0.005182 | 0.030457 | -9.7961 | 322.9553
92 | 537031.7964 | 0.005 | -0.0046184 | 0.0068066 | -42.7281 |  | 0.00019999 | -0.00018466 | 0.00019979 | -0.00018448 | 0.00019979 | -0.00018448 | 0.00019979 | -0.00018448 | -71.3106 | -42.7175 | 0.004995 | -0.0046138 | 0.0067998 | -42.7281 | -0.10002 | -0.9998 | -0.0001847 | -0.9988 | -0.00018451 | -0.9988 | -0.00018451 | -0.9988 | -0.00018451 | -0.010427 | -179.9894 | 0.030013 | -0.0046184 | 0.030366 | -8.7482 | 346.1231
93 | 575439.9373 | 0.005 | -0.0040769 | 0.0064514 | -39.1929 |  | 0.00019999 | -0.00016301 | 0.00019979 | -0.00016285 | 0.00019979 | -0.00016285 | 0.00019979 | -0.00016285 | -71.7761 | -39.1836 | 0.004995 | -0.0040728 | 0.006445 | -39.1929 | -0.10002 | -0.9998 | -0.00016304 | -0.9988 | -0.00016288 | -0.9988 | -0.00016288 | -0.9988 | -0.00016288 | -0.010427 | -179.9907 | 0.030013 | -0.0040769 | 0.030288 | -7.7357 | 369.4797
94 | 616595.0019 | 0.005 | -0.0035548 | 0.0061349 | -35.4113 |  | 0.00019998 | -0.00014214 | 0.00019978 | -0.00014199 | 0.00019978 | -0.00014199 | 0.00019978 | -0.00014199 | -72.2131 | -35.4032 | 0.004995 | -0.0035512 | 0.0061287 | -35.4113 | -0.10002 | -0.9998 | -0.00014216 | -0.9988 | -0.00014202 | -0.9988 | -0.00014202 | -0.9988 | -0.00014202 | -0.010427 | -179.9919 | 0.030013 | -0.0035548 | 0.030222 | -6.7549 | 392.6313
95 | 660693.448 | 0.005 | -0.0030497 | 0.0058567 | -31.3807 |  | 0.00019997 | -0.00012194 | 0.00019977 | -0.00012182 | 0.00019977 | -0.00012182 | 0.00019977 | -0.00012182 | -72.6162 | -31.3737 | 0.004995 | -0.0030467 | 0.0058508 | -31.3807 | -0.10002 | -0.9998 | -0.00012196 | -0.9988 | -0.00012184 | -0.9988 | -0.00012184 | -0.9988 | -0.00012184 | -0.010427 | -179.993 | 0.030013 | -0.0030497 | 0.030167 | -5.8022 | 415.0879
96 | 707945.7844 | 0.005 | -0.0025592 | 0.0056169 | -27.1049 |  | 0.00019997 | -0.00010233 | 0.00019977 | -0.00010222 | 0.00019977 | -0.00010222 | 0.00019977 | -0.00010222 | -72.9793 | -27.099 | 0.004995 | -0.0025566 | 0.0056113 | -27.1048 | -0.10002 | -0.9998 | -0.00010235 | -0.9988 | -0.00010224 | -0.9988 | -0.00010224 | -0.9988 | -0.00010224 | -0.010427 | -179.9941 | 0.030013 | -0.0025592 | 0.030121 | -4.8738 | 436.2662
97 | 758577.575 | 0.005 | -0.0020808 | 0.0054157 | -22.5955 |  | 0.00019997 | -8.32e-05 | 0.00019977 | -8.3117e-05 | 0.00019977 | -8.3117e-05 | 0.00019977 | -8.3117e-05 | -73.2961 | -22.5907 | 0.004995 | -0.0020788 | 0.0054103 | -22.5955 | -0.10002 | -0.9998 | -8.3217e-05 | -0.9988 | -8.3134e-05 | -0.9988 | -8.3134e-05 | -0.9988 | -8.3134e-05 | -0.010427 | -179.9952 | 0.030013 | -0.0020808 | 0.030085 | -3.9661 | 455.5055
98 | 812830.5162 | 0.005 | -0.0016124 | 0.0052536 | -17.874 |  | 0.00019996 | -6.4472e-05 | 0.00019976 | -6.4408e-05 | 0.00019976 | -6.4408e-05 | 0.00019976 | -6.4408e-05 | -73.5601 | -17.8703 | 0.004995 | -0.0016108 | 0.0052483 | -17.874 | -0.10002 | -0.9998 | -6.4485e-05 | -0.9988 | -6.442e-05 | -0.9988 | -6.442e-05 | -0.9988 | -6.442e-05 | -0.010427 | -179.9963 | 0.030013 | -0.0016124 | 0.030056 | -3.0753 | 472.1023
99 | 870963.59 | 0.005 | -0.0011518 | 0.0051309 | -12.9718 |  | 0.00019996 | -4.6052e-05 | 0.00019976 | -4.6006e-05 | 0.00019976 | -4.6006e-05 | 0.00019976 | -4.6006e-05 | -73.7653 | -12.9692 | 0.004995 | -0.0011506 | 0.0051258 | -12.9718 | -0.10002 | -0.9998 | -4.6061e-05 | -0.9988 | -4.6015e-05 | -0.9988 | -4.6015e-05 | -0.9988 | -4.6015e-05 | -0.010427 | -179.9974 | 0.030013 | -0.0011518 | 0.030035 | -2.1977 | 485.3626
100 | 933254.3008 | 0.005 | -0.00069655 | 0.0050483 | -7.9309 |  | 0.00019996 | -2.7851e-05 | 0.00019976 | -2.7823e-05 | 0.00019976 | -2.7823e-05 | 0.00019976 | -2.7823e-05 | -73.9063 | -7.9293 | 0.004995 | -0.00069586 | 0.0050432 | -7.9309 | -0.10002 | -0.9998 | -2.7857e-05 | -0.9988 | -2.7829e-05 | -0.9988 | -2.7829e-05 | -0.9988 | -2.7829e-05 | -0.010427 | -179.9984 | 0.030013 | -0.00069655 | 0.030021 | -1.3295 | 494.6689
101 | 1000000 | 0.005 | -0.00024468 | 0.005006 | -2.8016 |  | 0.00019996 | -9.7834e-06 | 0.00019976 | -9.7736e-06 | 0.00019976 | -9.7736e-06 | 0.00019976 | -9.7736e-06 | -73.9794 | -2.8011 | 0.004995 | -0.00024444 | 0.005001 | -2.8016 | -0.10002 | -0.9998 | -9.7854e-06 | -0.9988 | -9.7756e-06 | -0.9988 | -9.7756e-06 | -0.9988 | -9.7756e-06 | -0.010427 | -179.9994 | 0.030013 | -0.00024468 | 0.030014 | -0.4671 | 499.5526
102 | 1071519.3052 | 0.005 | 0.00020602 | 0.0050042 | 2.3595 |  | 0.00019996 | 8.2375e-06 | 0.00019976 | 8.2293e-06 | 0.00019976 | 8.2293e-06 | 0.00019976 | 8.2293e-06 | -73.9825 | 2.359 | 0.004995 | 0.00020581 | 0.0049992 | 2.3595 | -0.10002 | -0.9998 | 8.2392e-06 | -0.9988 | 8.231e-06 | -0.9988 | 8.231e-06 | -0.9988 | 8.231e-06 | -0.010427 | 179.9995 | 0.030013 | 0.00020602 | 0.030013 | 0.3933 | 499.7553
103 | 1148153.6215 | 0.005 | 0.00065771 | 0.0050431 | 7.4938 |  | 0.00019996 | 2.6298e-05 | 0.00019976 | 2.6272e-05 | 0.00019976 | 2.6272e-05 | 0.00019976 | 2.6272e-05 | -73.9153 | 7.4923 | 0.004995 | 0.00065705 | 0.005038 | 7.4938 | -0.10002 | -0.9998 | 2.6303e-05 | -0.9988 | 2.6277e-05 | -0.9988 | 2.6277e-05 | -0.9988 | 2.6277e-05 | -0.010427 | 179.9985 | 0.030013 | 0.00065771 | 0.03002 | 1.2554 | 495.2663
104 | 1230268.7708 | 0.005 | 0.0011125 | 0.0051223 | 12.5444 |  | 0.00019996 | 4.4484e-05 | 0.00019976 | 4.4439e-05 | 0.00019976 | 4.4439e-05 | 0.00019976 | 4.4439e-05 | -73.78 | 12.5418 | 0.004995 | 0.0011114 | 0.0051172 | 12.5444 | -0.10002 | -0.9998 | 4.4493e-05 | -0.9988 | 4.4448e-05 | -0.9988 | 4.4448e-05 | -0.9988 | 4.4448e-05 | -0.010427 | 179.9975 | 0.030013 | 0.0011125 | 0.030033 | 2.1229 | 486.3234
105 | 1318256.7386 | 0.005 | 0.0015727 | 0.0052415 | 17.4602 |  | 0.00019996 | 6.2882e-05 | 0.00019976 | 6.2819e-05 | 0.00019976 | 6.2819e-05 | 0.00019976 | 6.2819e-05 | -73.5801 | 17.4566 | 0.004995 | 0.0015711 | 0.0052363 | 17.4602 | -0.10002 | -0.9998 | 6.2894e-05 | -0.9988 | 6.2831e-05 | -0.9988 | 6.2831e-05 | -0.9988 | 6.2831e-05 | -0.010427 | 179.9964 | 0.030013 | 0.0015727 | 0.030054 | 2.9996 | 473.3796
106 | 1412537.5446 | 0.005 | 0.0020403 | 0.0054003 | 22.1986 |  | 0.00019997 | 8.158e-05 | 0.00019977 | 8.1499e-05 | 0.00019977 | 8.1499e-05 | 0.00019977 | 8.1499e-05 | -73.3209 | 22.1939 | 0.004995 | 0.0020383 | 0.0053949 | 22.1986 | -0.10002 | -0.9998 | 8.1596e-05 | -0.9988 | 8.1515e-05 | -0.9988 | 8.1515e-05 | -0.9988 | 8.1515e-05 | -0.010427 | 179.9953 | 0.030013 | 0.0020403 | 0.030082 | 3.8891 | 457.0423
107 | 1513561.2484 | 0.005 | 0.0025177 | 0.0055981 | 26.7271 |  | 0.00019997 | 0.00010067 | 0.00019977 | 0.00010057 | 0.00019977 | 0.00010057 | 0.00019977 | 0.00010057 | -73.0084 | 26.7213 | 0.004995 | 0.0025152 | 0.0055925 | 26.7271 | -0.10002 | -0.9998 | 0.00010069 | -0.9988 | 0.00010059 | -0.9988 | 0.00010059 | -0.9988 | 0.00010059 | -0.010427 | 179.9942 | 0.030013 | 0.0025177 | 0.030118 | 4.7952 | 438.0017
108 | 1621810.0974 | 0.005 | 0.0030071 | 0.0058346 | 31.0236 |  | 0.00019997 | 0.00012024 | 0.00019977 | 0.00012012 | 0.00019977 | 0.00012012 | 0.00019977 | 0.00012012 | -72.649 | 31.0167 | 0.004995 | 0.0030041 | 0.0058288 | 31.0236 | -0.10002 | -0.9998 | 0.00012026 | -0.9988 | 0.00012014 | -0.9988 | 0.00012014 | -0.9988 | 0.00012014 | -0.010427 | 179.9931 | 0.030013 | 0.0030071 | 0.030163 | 5.7217 | 416.9629
109 | 1737800.8287 | 0.005 | 0.0035109 | 0.0061095 | 35.0755 |  | 0.00019998 | 0.00014038 | 0.00019978 | 0.00014024 | 0.00019978 | 0.00014024 | 0.00019978 | 0.00014024 | -72.2491 | 35.0674 | 0.004995 | 0.0035074 | 0.0061034 | 35.0755 | -0.10002 | -0.9998 | 0.00014041 | -0.9988 | 0.00014027 | -0.9988 | 0.00014027 | -0.9988 | 0.00014027 | -0.010427 | 179.992 | 0.030013 | 0.0035109 | 0.030217 | 6.6721 | 394.5918
110 | 1862087.1367 | 0.005 | 0.0040314 | 0.0064228 | 38.8784 |  | 0.00019999 | 0.00016119 | 0.00019979 | 0.00016103 | 0.00019979 | 0.00016103 | 0.00019979 | 0.00016103 | -71.8148 | 38.8692 | 0.004995 | 0.0040273 | 0.0064164 | 38.8784 | -0.10002 | -0.9998 | 0.00016122 | -0.9988 | 0.00016106 | -0.9988 | 0.00016106 | -0.9988 | 0.00016106 | -0.010427 | 179.9908 | 0.030013 | 0.0040314 | 0.030282 | 7.6504 | 371.4792
111 | 1995262.315 | 0.005 | 0.0045711 | 0.0067746 | 42.4345 |  | 0.00019999 | 0.00018277 | 0.00019979 | 0.00018259 | 0.00019979 | 0.00018259 | 0.00019979 | 0.00018259 | -71.3515 | 42.424 | 0.004995 | 0.0045666 | 0.0067678 | 42.4344 | -0.10002 | -0.9998 | 0.00018281 | -0.9988 | 0.00018263 | -0.9988 | 0.00018263 | -0.9988 | 0.00018263 | -0.010427 | 179.9895 | 0.030013 | 0.0045711 | 0.030359 | 8.66 | 348.1234
112 | 2137962.0895 | 0.005 | 0.0051327 | 0.0071655 | 45.7505 |  | 0.0002 | 0.00020523 | 0.0001998 | 0.00020502 | 0.0001998 | 0.00020502 | 0.0001998 | 0.00020502 | -70.8643 | 45.7387 | 0.004995 | 0.0051276 | 0.0071584 | 45.7504 | -0.10002 | -0.9998 | 0.00020527 | -0.9988 | 0.00020506 | -0.9988 | 0.00020506 | -0.9988 | 0.00020506 | -0.010427 | 179.9882 | 0.030013 | 0.0051327 | 0.030448 | 9.7048 | 324.9263
113 | 2290867.6528 | 0.005 | 0.0057188 | 0.0075964 | 48.8365 |  | 0.00020001 | 0.00022866 | 0.00019981 | 0.00022843 | 0.00019981 | 0.00022843 | 0.00019981 | 0.00022843 | -70.3571 | 48.8234 | 0.004995 | 0.0057131 | 0.0075888 | 48.8365 | -0.10002 | -0.9998 | 0.00022871 | -0.9988 | 0.00022848 | -0.9988 | 0.00022848 | -0.9988 | 0.00022848 | -0.010427 | 179.9869 | 0.030013 | 0.0057188 | 0.030552 | 10.7882 | 302.1988
114 | 2454708.9157 | 0.005 | 0.0063322 | 0.0080683 | 51.7048 |  | 0.00020002 | 0.00025319 | 0.00019982 | 0.00025293 | 0.00019982 | 0.00025293 | 0.00019982 | 0.00025293 | -69.8336 | 51.6903 | 0.004995 | 0.0063259 | 0.0080602 | 51.7048 | -0.10002 | -0.9998 | 0.00025324 | -0.9988 | 0.00025298 | -0.9988 | 0.00025298 | -0.9988 | 0.00025298 | -0.010427 | 179.9855 | 0.030013 | 0.0063322 | 0.030673 | 11.9138 | 280.172
115 | 2630267.9919 | 0.005 | 0.0069758 | 0.0085826 | 54.3684 |  | 0.00020004 | 0.00027892 | 0.00019984 | 0.00027864 | 0.00019984 | 0.00027864 | 0.00019984 | 0.00027864 | -69.2968 | 54.3525 | 0.004995 | 0.0069688 | 0.0085741 | 54.3684 | -0.10002 | -0.9998 | 0.00027898 | -0.9988 | 0.0002787 | -0.9988 | 0.0002787 | -0.9988 | 0.0002787 | -0.010427 | 179.984 | 0.030013 | 0.0069758 | 0.030813 | 13.0849 | 259.0096
116 | 2818382.9313 | 0.005 | 0.0076527 | 0.0091413 | 56.8409 |  | 0.00020005 | 0.00030599 | 0.00019985 | 0.00030568 | 0.00019985 | 0.00030568 | 0.00019985 | 0.00030568 | -68.749 | 56.8234 | 0.004995 | 0.0076451 | 0.0091322 | 56.8409 | -0.10002 | -0.9998 | 0.00030605 | -0.9988 | 0.00030574 | -0.9988 | 0.00030574 | -0.9988 | 0.00030574 | -0.010427 | 179.9825 | 0.030013 | 0.0076527 | 0.030973 | 14.3047 | 238.8213
117 | 3019951.7204 | 0.005 | 0.0083662 | 0.0097464 | 59.1355 |  | 0.00020007 | 0.00033451 | 0.00019987 | 0.00033418 | 0.00019987 | 0.00033418 | 0.00019987 | 0.00033418 | -68.1923 | 59.1164 | 0.004995 | 0.0083578 | 0.0097367 | 59.1355 | -0.10002 | -0.9998 | 0.00033458 | -0.9988 | 0.00033424 | -0.9988 | 0.00033424 | -0.9988 | 0.00033424 | -0.010427 | 179.9808 | 0.030013 | 0.0083662 | 0.031157 | 15.5761 | 219.674
118 | 3235936.5693 | 0.005 | 0.0091195 | 0.0104 | 61.2652 |  | 0.00020009 | 0.00036464 | 0.00019989 | 0.00036427 | 0.00019989 | 0.00036427 | 0.00019989 | 0.00036427 | -67.6283 | 61.2443 | 0.004995 | 0.0091104 | 0.01039 | 61.2652 | -0.10002 | -0.9998 | 0.00036471 | -0.9988 | 0.00036434 | -0.9988 | 0.00036434 | -0.9988 | 0.00036434 | -0.010427 | 179.9791 | 0.030013 | 0.0091195 | 0.031367 | 16.9018 | 201.6017
119 | 3467368.5045 | 0.005 | 0.0099164 | 0.011106 | 63.2422 |  | 0.00020012 | 0.0003965 | 0.00019992 | 0.0003961 | 0.00019992 | 0.0003961 | 0.00019992 | 0.0003961 | -67.0583 | 63.2194 | 0.004995 | 0.0099065 | 0.011095 | 63.2421 | -0.10002 | -0.9998 | 0.00039658 | -0.9988 | 0.00039618 | -0.9988 | 0.00039618 | -0.9988 | 0.00039618 | -0.010427 | 179.9773 | 0.030013 | 0.0099164 | 0.031608 | 18.2841 | 184.6143
120 | 3715352.291 | 0.005 | 0.010761 | 0.011866 | 65.0779 |  | 0.00020015 | 0.00043026 | 0.00019995 | 0.00042983 | 0.00019995 | 0.00042983 | 0.00019995 | 0.00042983 | -66.4834 | 65.0532 | 0.004995 | 0.01075 | 0.011854 | 65.0779 | -0.10002 | -0.9998 | 0.00043034 | -0.9988 | 0.00042991 | -0.9988 | 0.00042991 | -0.9988 | 0.00042991 | -0.010427 | 179.9753 | 0.030013 | 0.010761 | 0.031883 | 19.7248 | 168.7033
121 | 3981071.7055 | 0.005 | 0.011656 | 0.012683 | 66.783 |  | 0.00020018 | 0.00046607 | 0.00019998 | 0.0004656 | 0.00019998 | 0.0004656 | 0.00019998 | 0.0004656 | -65.9045 | 66.7563 | 0.004995 | 0.011645 | 0.012671 | 66.783 | -0.10002 | -0.9998 | 0.00046616 | -0.9988 | 0.00046569 | -0.9988 | 0.00046569 | -0.9988 | 0.00046569 | -0.010427 | 179.9733 | 0.030013 | 0.011656 | 0.032197 | 21.2253 | 153.8474
122 | 4265795.188 | 0.005 | 0.012608 | 0.013563 | 68.3674 |  | 0.00020021 | 0.0005041 | 0.00020001 | 0.0005036 | 0.00020001 | 0.0005036 | 0.00020001 | 0.0005036 | -65.3222 | 68.3385 | 0.004995 | 0.012595 | 0.013549 | 68.3673 | -0.10002 | -0.9998 | 0.0005042 | -0.9988 | 0.0005037 | -0.9988 | 0.0005037 | -0.9988 | 0.0005037 | -0.010427 | 179.9711 | 0.030013 | 0.012608 | 0.032553 | 22.7862 | 140.0164
123 | 4570881.8961 | 0.005 | 0.013619 | 0.014508 | 69.8401 |  | 0.00020026 | 0.00054454 | 0.00020006 | 0.000544 | 0.00020006 | 0.000544 | 0.00020006 | 0.000544 | -64.7372 | 69.8089 | 0.004995 | 0.013605 | 0.014493 | 69.8401 | -0.10002 | -0.9998 | 0.00054465 | -0.9988 | 0.00054411 | -0.9988 | 0.00054411 | -0.9988 | 0.00054411 | -0.010427 | 179.9688 | 0.030013 | 0.013619 | 0.032958 | 24.4075 | 127.1735
124 | 4897788.1937 | 0.005 | 0.014695 | 0.015523 | 71.2096 |  | 0.00020031 | 0.00058758 | 0.00020011 | 0.000587 | 0.00020011 | 0.000587 | 0.00020011 | 0.000587 | -64.1498 | 71.1759 | 0.004995 | 0.014681 | 0.015507 | 71.2096 | -0.10002 | -0.9998 | 0.0005877 | -0.9988 | 0.00058711 | -0.9988 | 0.00058711 | -0.9988 | 0.00058711 | -0.010427 | 179.9663 | 0.030013 | 0.014695 | 0.033417 | 26.0884 | 115.2782
125 | 5248074.6025 | 0.005 | 0.015842 | 0.016612 | 72.4835 |  | 0.00020036 | 0.00063343 | 0.00020016 | 0.0006328 | 0.00020016 | 0.0006328 | 0.00020016 | 0.0006328 | -63.5606 | 72.4472 | 0.004995 | 0.015826 | 0.016596 | 72.4835 | -0.10002 | -0.9998 | 0.00063356 | -0.9988 | 0.00063292 | -0.9988 | 0.00063292 | -0.9988 | 0.00063292 | -0.010427 | 179.9637 | 0.030013 | 0.015842 | 0.033937 | 27.8273 | 104.2875
126 | 5623413.2519 | 0.005 | 0.017064 | 0.017782 | 73.6689 |  | 0.00020043 | 0.0006823 | 0.00020023 | 0.00068162 | 0.00020023 | 0.00068162 | 0.00020023 | 0.00068162 | -62.9697 | 73.6298 | 0.004995 | 0.017047 | 0.017764 | 73.6689 | -0.10002 | -0.9998 | 0.00068244 | -0.9988 | 0.00068175 | -0.9988 | 0.00068175 | -0.9988 | 0.00068175 | -0.010427 | 179.9609 | 0.030013 | 0.017064 | 0.034524 | 29.6214 | 94.157
127 | 6025595.8607 | 0.005 | 0.018368 | 0.019036 | 74.7723 |  | 0.0002005 | 0.00073443 | 0.0002003 | 0.00073369 | 0.0002003 | 0.00073369 | 0.0002003 | 0.00073369 | -62.3776 | 74.7302 | 0.004995 | 0.01835 | 0.019017 | 74.7723 | -0.10002 | -0.9998 | 0.00073457 | -0.9988 | 0.00073384 | -0.9988 | 0.00073384 | -0.9988 | 0.00073384 | -0.010427 | 179.9579 | 0.030013 | 0.018368 | 0.035187 | 31.4671 | 84.8417
128 | 6456542.2903 | 0.005 | 0.019759 | 0.020382 | 75.7997 |  | 0.00020058 | 0.00079006 | 0.00020038 | 0.00078927 | 0.00020038 | 0.00078927 | 0.00020038 | 0.00078927 | -61.7842 | 75.7544 | 0.004995 | 0.01974 | 0.020362 | 75.7996 | -0.10002 | -0.9998 | 0.00079022 | -0.9988 | 0.00078943 | -0.9988 | 0.00078943 | -0.9988 | 0.00078943 | -0.010427 | 179.9547 | 0.030013 | 0.019759 | 0.035933 | 33.3597 | 76.2965
129 | 6918309.7092 | 0.005 | 0.021245 | 0.021825 | 76.7565 |  | 0.00020068 | 0.00084946 | 0.00020048 | 0.00084861 | 0.00020048 | 0.00084861 | 0.00020048 | 0.00084861 | -61.19 | 76.7078 | 0.004995 | 0.021224 | 0.021804 | 76.7565 | -0.10002 | -0.9998 | 0.00084963 | -0.9988 | 0.00084878 | -0.9988 | 0.00084878 | -0.9988 | 0.00084878 | -0.010427 | 179.9513 | 0.030013 | 0.021245 | 0.036771 | 35.2936 | 68.4771
130 | 7413102.413 | 0.005 | 0.022832 | 0.023373 | 77.6478 |  | 0.00020079 | 0.00091292 | 0.00020059 | 0.00091201 | 0.00020059 | 0.00091201 | 0.00020059 | 0.00091201 | -60.5949 | 77.5955 | 0.004995 | 0.022809 | 0.02335 | 77.6478 | -0.10002 | -0.9998 | 0.0009131 | -0.9988 | 0.00091219 | -0.9988 | 0.00091219 | -0.9988 | 0.00091219 | -0.010427 | 179.9477 | 0.030013 | 0.022832 | 0.03771 | 37.2622 | 61.3393
131 | 7943282.3472 | 0.005 | 0.024528 | 0.025033 | 78.4783 |  | 0.00020092 | 0.00098074 | 0.00020072 | 0.00097976 | 0.00020072 | 0.00097976 | 0.00020072 | 0.00097976 | -59.9991 | 78.4221 | 0.004995 | 0.024504 | 0.025008 | 78.4783 | -0.10002 | -0.9998 | 0.00098093 | -0.9988 | 0.00097995 | -0.9988 | 0.00097995 | -0.9988 | 0.00097995 | -0.010427 | 179.9438 | 0.030013 | 0.024528 | 0.038761 | 39.258 | 54.8401
132 | 8511380.382 | 0.005 | 0.026341 | 0.026812 | 79.2523 |  | 0.00020107 | 0.0010532 | 0.00020087 | 0.0010522 | 0.00020087 | 0.0010522 | 0.00020087 | 0.0010522 | -59.4027 | 79.1919 | 0.004995 | 0.026315 | 0.026785 | 79.2522 | -0.10002 | -0.9998 | 0.0010534 | -0.9988 | 0.0010524 | -0.9988 | 0.0010524 | -0.9988 | 0.0010524 | -0.010427 | 179.9396 | 0.030013 | 0.026341 | 0.039933 | 41.2728 | 48.9372
133 | 9120108.3936 | 0.005 | 0.02828 | 0.028719 | 79.9737 |  | 0.00020124 | 0.0011308 | 0.00020104 | 0.0011296 | 0.00020104 | 0.0011296 | 0.00020104 | 0.0011296 | -58.8059 | 79.9089 | 0.004995 | 0.028252 | 0.02869 | 79.9736 | -0.10002 | -0.9998 | 0.001131 | -0.9988 | 0.0011299 | -0.9988 | 0.0011299 | -0.9988 | 0.0011299 | -0.010427 | 179.9352 | 0.030013 | 0.02828 | 0.041237 | 43.298 | 43.5897
134 | 9772372.2096 | 0.005 | 0.030354 | 0.030763 | 80.6462 |  | 0.00020143 | 0.0012137 | 0.00020123 | 0.0012125 | 0.00020123 | 0.0012125 | 0.00020123 | 0.0012125 | -58.2086 | 80.5766 | 0.004995 | 0.030324 | 0.030733 | 80.6461 | -0.10002 | -0.9998 | 0.0012139 | -0.9988 | 0.0012127 | -0.9988 | 0.0012127 | -0.9988 | 0.0012127 | -0.010427 | 179.9304 | 0.030013 | 0.030354 | 0.042686 | 45.3244 | 38.7575
135 | 10471285.4805 | 0.005 | 0.032573 | 0.032955 | 81.2732 |  | 0.00020166 | 0.0013024 | 0.00020145 | 0.0013011 | 0.00020145 | 0.0013011 | 0.00020145 | 0.0013011 | -57.6109 | 81.1985 | 0.004995 | 0.032541 | 0.032922 | 81.2731 | -0.10002 | -0.9998 | 0.0013027 | -0.9988 | 0.0013014 | -0.9988 | 0.0013014 | -0.9988 | 0.0013014 | -0.010427 | 179.9253 | 0.030013 | 0.032573 | 0.044292 | 47.3429 | 34.4021
136 | 11220184.543 | 0.005 | 0.034947 | 0.035303 | 81.8578 |  | 0.00020191 | 0.0013973 | 0.00020171 | 0.0013959 | 0.00020171 | 0.0013959 | 0.00020171 | 0.0013959 | -57.0129 | 81.7778 | 0.004995 | 0.034912 | 0.035268 | 81.8578 | -0.10002 | -0.9998 | 0.0013976 | -0.9988 | 0.0013962 | -0.9988 | 0.0013962 | -0.9988 | 0.0013962 | -0.010427 | 179.9199 | 0.030013 | 0.034947 | 0.046066 | 49.3443 | 30.4863
137 | 12022644.3462 | 0.005 | 0.037489 | 0.037821 | 82.4031 |  | 0.00020221 | 0.0014989 | 0.00020201 | 0.0014974 | 0.00020201 | 0.0014974 | 0.00020201 | 0.0014974 | -56.4147 | 82.3172 | 0.0049951 | 0.037451 | 0.037783 | 82.403 | -0.10002 | -0.9998 | 0.0014992 | -0.9988 | 0.0014977 | -0.9988 | 0.0014977 | -0.9988 | 0.0014977 | -0.010427 | 179.9141 | 0.030013 | 0.037489 | 0.048022 | 51.32 | 26.9742
138 | 12882495.5169 | 0.005 | 0.040209 | 0.040518 | 82.9116 |  | 0.00020255 | 0.0016077 | 0.00020234 | 0.0016061 | 0.00020234 | 0.0016061 | 0.00020234 | 0.0016061 | -55.8162 | 82.8195 | 0.0049951 | 0.040168 | 0.040478 | 82.9115 | -0.10002 | -0.9998 | 0.001608 | -0.9988 | 0.0016064 | -0.9988 | 0.0016064 | -0.9988 | 0.0016064 | -0.010427 | 179.9078 | 0.030013 | 0.040209 | 0.050175 | 53.2616 | 23.8317
139 | 13803842.646 | 0.005 | 0.043121 | 0.04341 | 83.3859 |  | 0.00020293 | 0.0017241 | 0.00020273 | 0.0017224 | 0.00020273 | 0.0017224 | 0.00020273 | 0.0017224 | -55.2175 | 83.2871 | 0.0049951 | 0.043078 | 0.043366 | 83.3858 | -0.10002 | -0.9998 | 0.0017245 | -0.9988 | 0.0017228 | -0.9988 | 0.0017228 | -0.9988 | 0.0017228 | -0.010427 | 179.9012 | 0.030013 | 0.043121 | 0.052537 | 55.1616 | 21.0263
140 | 14791083.8817 | 0.005 | 0.046239 | 0.046508 | 83.8283 |  | 0.00020338 | 0.0018488 | 0.00020318 | 0.001847 | 0.00020318 | 0.001847 | 0.00020318 | 0.001847 | -54.6187 | 83.7224 | 0.0049951 | 0.046192 | 0.046462 | 83.8282 | -0.10002 | -0.9998 | 0.0018492 | -0.9988 | 0.0018473 | -0.9988 | 0.0018473 | -0.9988 | 0.0018473 | -0.010427 | 179.894 | 0.030013 | 0.046239 | 0.055125 | 57.0132 | 18.5275
141 | 15848931.9246 | 0.005 | 0.049577 | 0.049829 | 84.241 |  | 0.00020389 | 0.0019823 | 0.00020369 | 0.0019803 | 0.00020369 | 0.0019803 | 0.00020369 | 0.0019803 | -54.0196 | 84.1274 | 0.0049951 | 0.049528 | 0.049779 | 84.2409 | -0.10002 | -0.9998 | 0.0019827 | -0.9988 | 0.0019807 | -0.9988 | 0.0019807 | -0.9988 | 0.0019807 | -0.010427 | 179.8864 | 0.030013 | 0.049577 | 0.057954 | 58.8106 | 16.3062
142 | 16982436.5246 | 0.005 | 0.053152 | 0.053387 | 84.6261 |  | 0.00020448 | 0.0021252 | 0.00020427 | 0.0021231 | 0.00020427 | 0.0021231 | 0.00020427 | 0.0021231 | -53.4205 | 84.5043 | 0.0049951 | 0.053099 | 0.053334 | 84.6259 | -0.10002 | -0.9998 | 0.0021257 | -0.9988 | 0.0021235 | -0.9988 | 0.0021235 | -0.9988 | 0.0021235 | -0.010427 | 179.8782 | 0.030013 | 0.053152 | 0.06104 | 60.5488 | 14.3355
143 | 18197008.5861 | 0.005 | 0.056981 | 0.0572 | 84.9853 |  | 0.00020515 | 0.0022783 | 0.00020495 | 0.0022761 | 0.00020495 | 0.0022761 | 0.00020495 | 0.0022761 | -52.8213 | 84.8547 | 0.0049951 | 0.056924 | 0.057143 | 84.9851 | -0.10002 | -0.9998 | 0.0022788 | -0.9988 | 0.0022765 | -0.9988 | 0.0022765 | -0.9988 | 0.0022765 | -0.010427 | 179.8694 | 0.030013 | 0.056981 | 0.064402 | 62.2239 | 12.5903
144 | 19498445.9976 | 0.005 | 0.061083 | 0.061287 | 85.3204 |  | 0.00020593 | 0.0024423 | 0.00020572 | 0.0024399 | 0.00020572 | 0.0024399 | 0.00020572 | 0.0024399 | -52.2219 | 85.1804 | 0.0049951 | 0.061021 | 0.061226 | 85.3203 | -0.10002 | -0.9998 | 0.0024428 | -0.9988 | 0.0024404 | -0.9988 | 0.0024404 | -0.9988 | 0.0024404 | -0.010427 | 179.86 | 0.030013 | 0.061082 | 0.068057 | 63.8331 | 11.0475
145 | 20892961.3085 | 0.005 | 0.065475 | 0.065666 | 85.6331 |  | 0.00020682 | 0.0026179 | 0.00020661 | 0.0026153 | 0.00020661 | 0.0026153 | 0.00020661 | 0.0026153 | -51.6225 | 85.4831 | 0.0049952 | 0.06541 | 0.0656 | 85.6329 | -0.10002 | -0.9998 | 0.0026185 | -0.9988 | 0.0026159 | -0.9988 | 0.0026159 | -0.9988 | 0.0026159 | -0.010427 | 179.8499 | 0.030013 | 0.065475 | 0.072026 | 65.3742 | 9.6858
146 | 22387211.3857 | 0.005 | 0.07018 | 0.070358 | 85.9248 |  | 0.00020784 | 0.0028061 | 0.00020763 | 0.0028033 | 0.00020763 | 0.0028033 | 0.00020763 | 0.0028033 | -51.023 | 85.764 | 0.0049952 | 0.07011 | 0.070288 | 85.9247 | -0.10002 | -0.9998 | 0.0028066 | -0.9988 | 0.0028038 | -0.9988 | 0.0028038 | -0.9988 | 0.0028038 | -0.010427 | 179.8392 | 0.030013 | 0.07018 | 0.076328 | 66.846 | 8.4855
147 | 23988329.1902 | 0.005 | 0.07522 | 0.075386 | 86.1971 |  | 0.00020901 | 0.0030076 | 0.0002088 | 0.0030046 | 0.0002088 | 0.0030046 | 0.0002088 | 0.0030046 | -50.4234 | 86.0247 | 0.0049952 | 0.075145 | 0.075311 | 86.1969 | -0.10002 | -0.9998 | 0.0030082 | -0.9988 | 0.0030052 | -0.9988 | 0.0030052 | -0.9988 | 0.0030052 | -0.010427 | 179.8276 | 0.030013 | 0.07522 | 0.080987 | 68.2482 | 7.4289
148 | 25703957.8277 | 0.005 | 0.08062 | 0.080775 | 86.4511 |  | 0.00021035 | 0.0032235 | 0.00021014 | 0.0032202 | 0.00021014 | 0.0032202 | 0.00021014 | 0.0032202 | -49.8238 | 86.2664 | 0.0049953 | 0.080539 | 0.080694 | 86.4509 | -0.10002 | -0.99979 | 0.0032241 | -0.9988 | 0.0032209 | -0.9988 | 0.0032209 | -0.9988 | 0.0032209 | -0.010427 | 179.8152 | 0.030013 | 0.08062 | 0.086025 | 69.581 | 6.5
149 | 27542287.0334 | 0.005 | 0.086404 | 0.086548 | 86.6881 |  | 0.0002119 | 0.0034547 | 0.00021169 | 0.0034513 | 0.00021169 | 0.0034513 | 0.00021169 | 0.0034513 | -49.2241 | 86.4901 | 0.0049953 | 0.086317 | 0.086462 | 86.6879 | -0.10002 | -0.99979 | 0.0034554 | -0.99879 | 0.003452 | -0.99879 | 0.003452 | -0.99879 | 0.003452 | -0.010427 | 179.802 | 0.030013 | 0.086404 | 0.091468 | 70.8452 | 5.6841
150 | 29512092.2667 | 0.005 | 0.0926 | 0.092735 | 86.9093 |  | 0.00021367 | 0.0037025 | 0.00021346 | 0.0036988 | 0.00021346 | 0.0036988 | 0.00021346 | 0.0036988 | -48.6244 | 86.6971 | 0.0049953 | 0.092508 | 0.092642 | 86.9091 | -0.10002 | -0.99979 | 0.0037033 | -0.99879 | 0.0036996 | -0.99879 | 0.0036996 | -0.99879 | 0.0036996 | -0.010427 | 179.7878 | 0.030013 | 0.0926 | 0.097342 | 72.042 | 4.9683
151 | 31622776.6017 | 0.005 | 0.099239 | 0.099365 | 87.1157 |  | 0.00021571 | 0.0039679 | 0.00021549 | 0.0039639 | 0.00021549 | 0.0039639 | 0.00021549 | 0.0039639 | -48.0247 | 86.8883 | 0.0049954 | 0.09914 | 0.099265 | 87.1154 | -0.10002 | -0.99979 | 0.0039687 | -0.99879 | 0.0039648 | -0.99879 | 0.0039648 | -0.99879 | 0.0039648 | -0.010427 | 179.7726 | 0.030013 | 0.099239 | 0.10368 | 73.1732 | 4.3407
152 | 33884415.6139 | 0.005 | 0.10635 | 0.10647 | 87.3083 |  | 0.00021805 | 0.0042523 | 0.00021783 | 0.004248 | 0.00021783 | 0.004248 | 0.00021783 | 0.004248 | -47.4249 | 87.0646 | 0.0049955 | 0.10624 | 0.10636 | 87.308 | -0.10002 | -0.99979 | 0.0042532 | -0.99879 | 0.0042489 | -0.99879 | 0.0042489 | -0.99879 | 0.0042489 | -0.010427 | 179.7563 | 0.030013 | 0.10635 | 0.1105 | 74.2408 | 3.791
153 | 36307805.477 | 0.005 | 0.11397 | 0.11408 | 87.488 |  | 0.00022073 | 0.0045569 | 0.00022051 | 0.0045524 | 0.00022051 | 0.0045524 | 0.00022051 | 0.0045524 | -46.8251 | 87.2269 | 0.0049955 | 0.11386 | 0.11397 | 87.4877 | -0.10002 | -0.99979 | 0.0045579 | -0.99879 | 0.0045533 | -0.99879 | 0.0045533 | -0.99879 | 0.0045533 | -0.010427 | 179.7388 | 0.030013 | 0.11397 | 0.11786 | 75.2469 | 3.3098
154 | 38904514.4994 | 0.005 | 0.12214 | 0.12224 | 87.6557 |  | 0.00022381 | 0.0048833 | 0.00022359 | 0.0048785 | 0.00022359 | 0.0048785 | 0.00022359 | 0.0048785 | -46.2252 | 87.3759 | 0.0049956 | 0.12201 | 0.12212 | 87.6554 | -0.10002 | -0.99979 | 0.0048844 | -0.99879 | 0.0048795 | -0.99879 | 0.0048795 | -0.99879 | 0.0048795 | -0.010427 | 179.7201 | 0.030013 | 0.12214 | 0.12577 | 76.1941 | 2.8888
155 | 41686938.347 | 0.005 | 0.13088 | 0.13098 | 87.8122 |  | 0.00022735 | 0.005233 | 0.00022712 | 0.0052278 | 0.00022712 | 0.0052278 | 0.00022712 | 0.0052278 | -45.6254 | 87.5123 | 0.0049957 | 0.13075 | 0.13085 | 87.8119 | -0.10002 | -0.99979 | 0.0052342 | -0.99879 | 0.005229 | -0.99879 | 0.005229 | -0.99879 | 0.005229 | -0.010427 | 179.7 | 0.030013 | 0.13088 | 0.13428 | 77.0847 | 2.5207
156 | 44668359.2151 | 0.005 | 0.14025 | 0.14034 | 87.9583 |  | 0.00023141 | 0.0056077 | 0.00023118 | 0.0056021 | 0.00023118 | 0.0056021 | 0.00023118 | 0.0056021 | -45.0255 | 87.6369 | 0.0049958 | 0.14011 | 0.1402 | 87.958 | -0.10002 | -0.99978 | 0.005609 | -0.99878 | 0.0056034 | -0.99878 | 0.0056034 | -0.99878 | 0.0056034 | -0.010427 | 179.6786 | 0.030013 | 0.14025 | 0.14343 | 77.9215 | 2.199
157 | 47863009.2323 | 0.005 | 0.1503 | 0.15038 | 88.0946 |  | 0.00023608 | 0.0060092 | 0.00023584 | 0.0060032 | 0.00023584 | 0.0060032 | 0.00023584 | 0.0060032 | -44.4257 | 87.7502 | 0.0049959 | 0.15014 | 0.15023 | 88.0943 | -0.10002 | -0.99978 | 0.0060106 | -0.99878 | 0.0060045 | -0.99878 | 0.0060045 | -0.99878 | 0.0060045 | -0.010427 | 179.6555 | 0.030013 | 0.1503 | 0.15326 | 78.707 | 1.9179
158 | 51286138.3991 | 0.005 | 0.16105 | 0.16113 | 88.2218 |  | 0.00024143 | 0.0064393 | 0.00024119 | 0.0064329 | 0.00024119 | 0.0064329 | 0.00024119 | 0.0064329 | -43.8258 | 87.8528 | 0.004996 | 0.16089 | 0.16097 | 88.2214 | -0.10002 | -0.99978 | 0.0064408 | -0.99878 | 0.0064344 | -0.99878 | 0.0064344 | -0.99878 | 0.0064344 | -0.010427 | 179.6309 | 0.030013 | 0.16105 | 0.16383 | 79.4439 | 1.6725
159 | 54954087.3858 | 0.005 | 0.17258 | 0.17265 | 88.3405 |  | 0.00024758 | 0.0069002 | 0.00024734 | 0.0068933 | 0.00024734 | 0.0068933 | 0.00024734 | 0.0068933 | -43.2259 | 87.9451 | 0.0049962 | 0.17241 | 0.17248 | 88.3401 | -0.10002 | -0.99978 | 0.0069018 | -0.99878 | 0.0068949 | -0.99878 | 0.0068949 | -0.99878 | 0.0068949 | -0.010427 | 179.6045 | 0.030013 | 0.17258 | 0.17517 | 80.1346 | 1.4583
160 | 58884365.5356 | 0.005 | 0.18493 | 0.185 | 88.4513 |  | 0.00025464 | 0.007394 | 0.00025439 | 0.0073866 | 0.00025439 | 0.0073866 | 0.00025439 | 0.0073866 | -42.626 | 88.0275 | 0.0049964 | 0.18475 | 0.18482 | 88.4509 | -0.10002 | -0.99977 | 0.0073957 | -0.99877 | 0.0073884 | -0.99877 | 0.0073884 | -0.99877 | 0.0073884 | -0.010427 | 179.5762 | 0.030013 | 0.18493 | 0.18735 | 80.7818 | 1.2713
161 | 63095734.448 | 0.005 | 0.19817 | 0.19823 | 88.5547 |  | 0.00026275 | 0.007923 | 0.00026249 | 0.0079151 | 0.00026249 | 0.0079151 | 0.00026249 | 0.0079151 | -42.0261 | 88.1006 | 0.0049966 | 0.19797 | 0.19803 | 88.5542 | -0.10002 | -0.99977 | 0.007925 | -0.99877 | 0.0079171 | -0.99877 | 0.0079171 | -0.99877 | 0.0079171 | -0.010427 | 179.5458 | 0.030013 | 0.19817 | 0.20043 | 81.3879 | 1.1082
162 | 67608297.5392 | 0.005 | 0.21235 | 0.21241 | 88.6511 |  | 0.00027206 | 0.0084899 | 0.00027179 | 0.0084814 | 0.00027179 | 0.0084814 | 0.00027179 | 0.0084814 | -41.4262 | 88.1646 | 0.0049968 | 0.21214 | 0.21219 | 88.6507 | -0.10002 | -0.99976 | 0.0084921 | -0.99876 | 0.0084836 | -0.99876 | 0.0084836 | -0.99876 | 0.0084836 | -0.010427 | 179.5133 | 0.030013 | 0.21235 | 0.21446 | 81.9552 | 0.96587
163 | 72443596.0075 | 0.005 | 0.22754 | 0.2276 | 88.7412 |  | 0.00028274 | 0.0090973 | 0.00028246 | 0.0090882 | 0.00028246 | 0.0090882 | 0.00028246 | 0.0090882 | -40.8263 | 88.2198 | 0.0049971 | 0.22731 | 0.22737 | 88.7407 | -0.10002 | -0.99976 | 0.0090997 | -0.99876 | 0.0090906 | -0.99876 | 0.0090906 | -0.99876 | 0.0090906 | -0.010427 | 179.4785 | 0.030013 | 0.22754 | 0.22951 | 82.486 | 0.84177
164 | 77624711.6629 | 0.005 | 0.24382 | 0.24387 | 88.8252 |  | 0.00029501 | 0.009748 | 0.00029472 | 0.0097383 | 0.00029472 | 0.0097383 | 0.00029472 | 0.0097383 | -40.2264 | 88.2665 | 0.0049974 | 0.24358 | 0.24363 | 88.8247 | -0.10002 | -0.99975 | 0.0097507 | -0.99875 | 0.0097409 | -0.99875 | 0.0097409 | -0.99875 | 0.0097409 | -0.010427 | 179.4412 | 0.030013 | 0.24382 | 0.24566 | 82.9825 | 0.73355
165 | 83176377.1103 | 0.005 | 0.26127 | 0.26131 | 88.9036 |  | 0.0003091 | 0.010445 | 0.00030879 | 0.010435 | 0.00030879 | 0.010435 | 0.00030879 | 0.010435 | -39.6265 | 88.305 | 0.0049977 | 0.261 | 0.26105 | 88.903 | -0.10002 | -0.99975 | 0.010448 | -0.99875 | 0.010438 | -0.99875 | 0.010438 | -0.99875 | 0.010438 | -0.010427 | 179.4012 | 0.030013 | 0.26127 | 0.26298 | 83.4468 | 0.6392
166 | 89125093.8134 | 0.005 | 0.27996 | 0.28 | 88.9768 |  | 0.00032527 | 0.011192 | 0.00032495 | 0.011181 | 0.00032495 | 0.011181 | 0.00032495 | 0.011181 | -39.0266 | 88.3354 | 0.0049981 | 0.27968 | 0.27972 | 88.9762 | -0.10002 | -0.99974 | 0.011196 | -0.99874 | 0.011184 | -0.99874 | 0.011184 | -0.99874 | 0.011184 | -0.010427 | 179.3584 | 0.030013 | 0.27996 | 0.28156 | 83.8809 | 0.55695
167 | 95499258.6021 | 0.005 | 0.29998 | 0.30003 | 89.0451 |  | 0.00034384 | 0.011993 | 0.00034349 | 0.011981 | 0.00034349 | 0.011981 | 0.00034349 | 0.011981 | -38.4267 | 88.3578 | 0.0049986 | 0.29968 | 0.29973 | 89.0444 | -0.10002 | -0.99973 | 0.011997 | -0.99873 | 0.011985 | -0.99873 | 0.011985 | -0.99873 | 0.011985 | -0.010427 | 179.3125 | 0.030013 | 0.29998 | 0.30148 | 84.2866 | 0.48526
168 | 102329299.2281 | 0.005 | 0.32144 | 0.32148 | 89.1088 |  | 0.00036516 | 0.01285 | 0.00036479 | 0.012838 | 0.00036479 | 0.012838 | 0.00036479 | 0.012838 | -37.8268 | 88.3723 | 0.0049991 | 0.32112 | 0.32116 | 89.1081 | -0.10002 | -0.99972 | 0.012855 | -0.99872 | 0.012842 | -0.99872 | 0.012842 | -0.99872 | 0.012842 | -0.010427 | 179.2633 | 0.030014 | 0.32144 | 0.32284 | 84.6657 | 0.42278
169 | 109647819.6143 | 0.005 | 0.34444 | 0.34447 | 89.1683 |  | 0.00038963 | 0.013769 | 0.00038924 | 0.013756 | 0.00038924 | 0.013756 | 0.00038924 | 0.013756 | -37.2269 | 88.3791 | 0.0049997 | 0.34409 | 0.34413 | 89.1675 | -0.10002 | -0.99971 | 0.013774 | -0.99871 | 0.01376 | -0.99871 | 0.01376 | -0.99871 | 0.01376 | -0.010427 | 179.2106 | 0.030014 | 0.34444 | 0.34574 | 85.0199 | 0.36833
170 | 117489755.494 | 0.005 | 0.36908 | 0.36911 | 89.2238 |  | 0.00041773 | 0.014754 | 0.00041731 | 0.014739 | 0.00041731 | 0.014739 | 0.00041731 | 0.014739 | -36.6271 | 88.3782 | 0.0050004 | 0.36871 | 0.36874 | 89.223 | -0.10002 | -0.99969 | 0.014759 | -0.99869 | 0.014745 | -0.99869 | 0.014745 | -0.99869 | 0.014745 | -0.010427 | 179.1542 | 0.030014 | 0.36908 | 0.37029 | 85.3508 | 0.32087
171 | 125892541.1794 | 0.005 | 0.39548 | 0.39551 | 89.2756 |  | 0.00044999 | 0.015809 | 0.00044954 | 0.015793 | 0.00044954 | 0.015793 | 0.00044954 | 0.015793 | -36.0272 | 88.3695 | 0.0050012 | 0.39508 | 0.39511 | 89.2747 | -0.10002 | -0.99967 | 0.015815 | -0.99868 | 0.015799 | -0.99868 | 0.015799 | -0.99868 | 0.015799 | -0.010427 | 179.0937 | 0.030014 | 0.39548 | 0.39661 | 85.6599 | 0.27953
172 | 134896288.2592 | 0.005 | 0.42376 | 0.42379 | 89.324 |  | 0.00048703 | 0.016939 | 0.00048654 | 0.016922 | 0.00048654 | 0.016922 | 0.00048654 | 0.016922 | -35.4274 | 88.3531 | 0.0050022 | 0.42334 | 0.42337 | 89.323 | -0.10002 | -0.99966 | 0.016946 | -0.99866 | 0.016929 | -0.99866 | 0.016929 | -0.99866 | 0.016929 | -0.010427 | 179.0288 | 0.030014 | 0.42376 | 0.42483 | 85.9486 | 0.2435
173 | 144543977.0746 | 0.005 | 0.45407 | 0.4541 | 89.3691 |  | 0.00052955 | 0.01815 | 0.00052902 | 0.018132 | 0.00052902 | 0.018132 | 0.00052902 | 0.018132 | -34.8276 | 88.3288 | 0.0050032 | 0.45362 | 0.45365 | 89.3681 | -0.10002 | -0.99964 | 0.018158 | -0.99864 | 0.01814 | -0.99864 | 0.01814 | -0.99864 | 0.01814 | -0.010427 | 178.9594 | 0.030015 | 0.45407 | 0.45507 | 86.2182 | 0.21212
174 | 154881661.8912 | 0.005 | 0.48655 | 0.48658 | 89.4112 |  | 0.00057836 | 0.019447 | 0.00057779 | 0.019428 | 0.00057779 | 0.019428 | 0.00057779 | 0.019428 | -34.2278 | 88.2965 | 0.0050045 | 0.48607 | 0.48609 | 89.4101 | -0.10002 | -0.99961 | 0.019456 | -0.99861 | 0.019437 | -0.99861 | 0.019437 | -0.99861 | 0.019437 | -0.010427 | 178.8849 | 0.030015 | 0.48655 | 0.48748 | 86.47 | 0.18477
175 | 165958690.7438 | 0.005 | 0.52135 | 0.52138 | 89.4505 |  | 0.00063441 | 0.020837 | 0.00063377 | 0.020816 | 0.00063377 | 0.020816 | 0.00063377 | 0.020816 | -33.6281 | 88.2561 | 0.0050059 | 0.52083 | 0.52086 | 89.4493 | -0.10002 | -0.99958 | 0.020848 | -0.99858 | 0.020827 | -0.99858 | 0.020827 | -0.99858 | 0.020827 | -0.010427 | 178.8052 | 0.030015 | 0.52135 | 0.52222 | 86.705 | 0.16095
176 | 177827941.0039 | 0.005 | 0.55864 | 0.55867 | 89.4872 |  | 0.00069874 | 0.022326 | 0.00069805 | 0.022303 | 0.00069805 | 0.022303 | 0.00069805 | 0.022303 | -33.0283 | 88.2074 | 0.0050075 | 0.55809 | 0.55811 | 89.4859 | -0.10002 | -0.99955 | 0.022338 | -0.99855 | 0.022316 | -0.99855 | 0.022316 | -0.99855 | 0.022316 | -0.010427 | 178.7197 | 0.030016 | 0.55864 | 0.55945 | 86.9245 | 0.14019
177 | 190546071.7963 | 0.005 | 0.5986 | 0.59862 | 89.5214 |  | 0.0007726 | 0.023921 | 0.00077183 | 0.023897 | 0.00077183 | 0.023897 | 0.00077183 | 0.023897 | -32.4287 | 88.1501 | 0.0050093 | 0.598 | 0.59802 | 89.5201 | -0.10002 | -0.99951 | 0.023936 | -0.99851 | 0.023912 | -0.99851 | 0.023912 | -0.99851 | 0.023912 | -0.010427 | 178.6282 | 0.030016 | 0.5986 | 0.59935 | 87.1294 | 0.12211
178 | 204173794.467 | 0.005 | 0.64141 | 0.64143 | 89.5534 |  | 0.00085739 | 0.025629 | 0.00085653 | 0.025604 | 0.00085653 | 0.025604 | 0.00085653 | 0.025604 | -31.829 | 88.084 | 0.0050114 | 0.64077 | 0.64079 | 89.5519 | -0.10002 | -0.99947 | 0.025647 | -0.99847 | 0.025622 | -0.99847 | 0.025622 | -0.99847 | 0.025622 | -0.010427 | 178.5301 | 0.030017 | 0.64141 | 0.64212 | 87.3207 | 0.10636
179 | 218776162.395 | 0.005 | 0.68729 | 0.68731 | 89.5832 |  | 0.00095472 | 0.02746 | 0.00095377 | 0.027432 | 0.00095377 | 0.027432 | 0.00095377 | 0.027432 | -31.2295 | 88.0087 | 0.0050139 | 0.6866 | 0.68662 | 89.5816 | -0.10002 | -0.99942 | 0.027481 | -0.99842 | 0.027453 | -0.99842 | 0.027453 | -0.99842 | 0.027453 | -0.010427 | 178.4249 | 0.030017 | 0.68729 | 0.68795 | 87.4992 | 0.092645
180 | 234422881.532 | 0.005 | 0.73645 | 0.73646 | 89.611 |  | 0.0010665 | 0.029421 | 0.0010654 | 0.029391 | 0.0010654 | 0.029391 | 0.0010654 | 0.029391 | -30.63 | 87.924 | 0.0050167 | 0.73571 | 0.73573 | 89.6093 | -0.10002 | -0.99937 | 0.029446 | -0.99837 | 0.029416 | -0.99837 | 0.029416 | -0.99837 | 0.029416 | -0.010427 | 178.3123 | 0.030018 | 0.73645 | 0.73706 | 87.6659 | 0.080695
181 | 251188643.151 | 0.005 | 0.78912 | 0.78913 | 89.637 |  | 0.0011947 | 0.031521 | 0.0011935 | 0.031489 | 0.0011935 | 0.031489 | 0.0011935 | 0.031489 | -30.0305 | 87.8294 | 0.0050199 | 0.78833 | 0.78835 | 89.6352 | -0.10002 | -0.9993 | 0.031551 | -0.9983 | 0.031519 | -0.9983 | 0.031519 | -0.9983 | 0.031519 | -0.010427 | 178.1916 | 0.030019 | 0.78912 | 0.78969 | 87.8215 | 0.070285
182 | 269153480.3927 | 0.005 | 0.84556 | 0.84557 | 89.6612 |  | 0.0013419 | 0.03377 | 0.0013406 | 0.033736 | 0.0013406 | 0.033736 | 0.0013406 | 0.033736 | -29.4312 | 87.7244 | 0.0050236 | 0.84471 | 0.84473 | 89.6593 | -0.10002 | -0.99923 | 0.033806 | -0.99823 | 0.033772 | -0.99823 | 0.033772 | -0.99823 | 0.033772 | -0.010427 | 178.0623 | 0.03002 | 0.84556 | 0.84609 | 87.9667 | 0.061218
183 | 288403150.3127 | 0.005 | 0.90603 | 0.90605 | 89.6838 |  | 0.0015109 | 0.036179 | 0.0015094 | 0.036143 | 0.0015094 | 0.036143 | 0.0015094 | 0.036143 | -28.8319 | 87.6087 | 0.0050278 | 0.90513 | 0.90514 | 89.6817 | -0.10002 | -0.99914 | 0.036222 | -0.99814 | 0.036186 | -0.99814 | 0.036186 | -0.99814 | 0.036186 | -0.010427 | 177.9238 | 0.030021 | 0.90603 | 0.90653 | 88.1022 | 0.05332
184 | 309029543.2514 | 0.005 | 0.97083 | 0.97085 | 89.7049 |  | 0.0017048 | 0.038759 | 0.0017031 | 0.038721 | 0.0017031 | 0.038721 | 0.0017031 | 0.038721 | -28.2328 | 87.4815 | 0.0050327 | 0.96986 | 0.96988 | 89.7027 | -0.10002 | -0.99905 | 0.038811 | -0.99805 | 0.038772 | -0.99805 | 0.038772 | -0.99805 | 0.038772 | -0.010427 | 177.7753 | 0.030022 | 0.97083 | 0.9713 | 88.2288 | 0.046441
185 | 331131121.4826 | 0.005 | 1.0403 | 1.0403 | 89.7246 |  | 0.0019274 | 0.041522 | 0.0019255 | 0.041481 | 0.0019255 | 0.041481 | 0.0019255 | 0.041481 | -27.6337 | 87.3423 | 0.0050382 | 1.0392 | 1.0392 | 89.7222 | -0.10002 | -0.99893 | 0.041584 | -0.99794 | 0.041543 | -0.99794 | 0.041543 | -0.99794 | 0.041543 | -0.010427 | 177.6162 | 0.030023 | 1.0403 | 1.0407 | 88.3468 | 0.040449
186 | 354813389.2336 | 0.005 | 1.1147 | 1.1147 | 89.743 |  | 0.0021828 | 0.044481 | 0.0021806 | 0.044436 | 0.0021806 | 0.044436 | 0.0021806 | 0.044436 | -27.0348 | 87.1906 | 0.0050446 | 1.1136 | 1.1136 | 89.7404 | -0.10002 | -0.99881 | 0.044556 | -0.99781 | 0.044511 | -0.99781 | 0.044511 | -0.99781 | 0.044511 | -0.010427 | 177.4458 | 0.030025 | 1.1147 | 1.1151 | 88.457 | 0.03523
187 | 380189396.3206 | 0.005 | 1.1944 | 1.1944 | 89.7601 |  | 0.0024759 | 0.047648 | 0.0024734 | 0.0476 | 0.0024734 | 0.0476 | 0.0024734 | 0.0476 | -26.4361 | 87.0254 | 0.005052 | 1.1932 | 1.1932 | 89.7574 | -0.10002 | -0.99866 | 0.047739 | -0.99766 | 0.047691 | -0.99766 | 0.047691 | -0.99766 | 0.047691 | -0.010426 | 177.2632 | 0.030027 | 1.1944 | 1.1948 | 88.5599 | 0.030684
188 | 407380277.8041 | 0.005 | 1.2798 | 1.2798 | 89.7762 |  | 0.0028122 | 0.051038 | 0.0028094 | 0.050987 | 0.0028094 | 0.050987 | 0.0028094 | 0.050987 | -25.8376 | 86.8462 | 0.0050605 | 1.2785 | 1.2785 | 89.7732 | -0.10002 | -0.99849 | 0.051149 | -0.99749 | 0.051098 | -0.99749 | 0.051098 | -0.99749 | 0.051098 | -0.010426 | 177.0675 | 0.030029 | 1.2798 | 1.2802 | 88.6559 | 0.026725
189 | 436515832.2402 | 0.005 | 1.3713 | 1.3714 | 89.7911 |  | 0.0031981 | 0.054668 | 0.0031949 | 0.054613 | 0.0031949 | 0.054613 | 0.0031949 | 0.054613 | -25.2393 | 86.652 | 0.0050701 | 1.37 | 1.37 | 89.788 | -0.10002 | -0.9983 | 0.054802 | -0.9973 | 0.054747 | -0.9973 | 0.054747 | -0.9973 | 0.054747 | -0.010426 | 176.8579 | 0.030031 | 1.3713 | 1.3717 | 88.7455 | 0.023276
190 | 467735141.2872 | 0.005 | 1.4694 | 1.4694 | 89.805 |  | 0.0036407 | 0.058551 | 0.0036371 | 0.058493 | 0.0036371 | 0.058493 | 0.0036371 | 0.058493 | -24.6412 | 86.4419 | 0.0050813 | 1.468 | 1.468 | 89.8017 | -0.10002 | -0.99807 | 0.058715 | -0.99708 | 0.058656 | -0.99708 | 0.058656 | -0.99708 | 0.058656 | -0.010426 | 176.6333 | 0.030034 | 1.4694 | 1.4697 | 88.8291 | 0.020272
191 | 501187233.6273 | 0.005 | 1.5745 | 1.5745 | 89.8181 |  | 0.0041485 | 0.062707 | 0.0041444 | 0.062644 | 0.0041444 | 0.062644 | 0.0041444 | 0.062644 | -24.0434 | 86.215 | 0.0050941 | 1.5729 | 1.573 | 89.8144 | -0.10002 | -0.99782 | 0.062906 | -0.99682 | 0.062843 | -0.99682 | 0.062843 | -0.99682 | 0.062843 | -0.010426 | 176.3927 | 0.030037 | 1.5745 | 1.5748 | 88.9071 | 0.017656
192 | 537031796.3703 | 0.005 | 1.6871 | 1.6871 | 89.8302 |  | 0.0047308 | 0.067152 | 0.0047261 | 0.067085 | 0.0047261 | 0.067085 | 0.0047261 | 0.067085 | -23.446 | 85.9702 | 0.0051087 | 1.6854 | 1.6854 | 89.8263 | -0.10002 | -0.99753 | 0.067395 | -0.99653 | 0.067328 | -0.99653 | 0.067328 | -0.99653 | 0.067328 | -0.010425 | 176.1349 | 0.030041 | 1.6871 | 1.6874 | 88.9799 | 0.015377
193 | 575439937.3372 | 0.005 | 1.8078 | 1.8078 | 89.8415 |  | 0.0053986 | 0.071907 | 0.0053932 | 0.071835 | 0.0053932 | 0.071835 | 0.0053932 | 0.071835 | -22.8489 | 85.7064 | 0.0051256 | 1.806 | 1.806 | 89.8374 | -0.10002 | -0.99719 | 0.072203 | -0.99619 | 0.072131 | -0.99619 | 0.072131 | -0.99619 | 0.072131 | -0.010425 | 175.8586 | 0.030045 | 1.8078 | 1.808 | 89.0478 | 0.013392
194 | 616595001.8615 | 0.005 | 1.9371 | 1.9371 | 89.8521 |  | 0.0061642 | 0.07699 | 0.0061581 | 0.076913 | 0.0061581 | 0.076913 | 0.0061581 | 0.076913 | -22.2522 | 85.4224 | 0.0051449 | 1.9351 | 1.9352 | 89.8477 | -0.10002 | -0.9968 | 0.077352 | -0.99581 | 0.077274 | -0.99581 | 0.077274 | -0.99581 | 0.077274 | -0.010425 | 175.5627 | 0.03005 | 1.9371 | 1.9373 | 89.1112 | 0.011664
195 | 660693448.0076 | 0.005 | 2.0756 | 2.0756 | 89.862 |  | 0.0070418 | 0.082424 | 0.0070348 | 0.082341 | 0.0070348 | 0.082341 | 0.0070348 | 0.082341 | -21.656 | 85.1168 | 0.0051672 | 2.0735 | 2.0736 | 89.8572 | -0.10002 | -0.99636 | 0.082866 | -0.99536 | 0.082783 | -0.99536 | 0.082783 | -0.99536 | 0.082783 | -0.010424 | 175.2457 | 0.030056 | 2.0756 | 2.0758 | 89.1704 | 0.010158
196 | 707945784.3841 | 0.005 | 2.2241 | 2.2241 | 89.8712 |  | 0.0080475 | 0.088229 | 0.0080395 | 0.088141 | 0.0080395 | 0.088141 | 0.0080395 | 0.088141 | -21.0604 | 84.7884 | 0.0051927 | 2.2218 | 2.2219 | 89.8661 | -0.10002 | -0.99585 | 0.08877 | -0.99486 | 0.088681 | -0.99486 | 0.088681 | -0.99486 | 0.088681 | -0.010424 | 174.9062 | 0.030062 | 2.2241 | 2.2243 | 89.2256 | 0.0088468
197 | 758577575.0292 | 0.005 | 2.3831 | 2.3831 | 89.8798 |  | 0.0091997 | 0.09443 | 0.0091905 | 0.094335 | 0.0091905 | 0.094335 | 0.0091905 | 0.094335 | -20.4655 | 84.4356 | 0.0052219 | 2.3808 | 2.3808 | 89.8743 | -0.10002 | -0.99527 | 0.09509 | -0.99427 | 0.094995 | -0.99427 | 0.094995 | -0.99427 | 0.094995 | -0.010423 | 174.5424 | 0.030069 | 2.3831 | 2.3833 | 89.2771 | 0.0077046
198 | 812830516.1641 | 0.005 | 2.5536 | 2.5536 | 89.8878 |  | 0.010519 | 0.10105 | 0.010509 | 0.10095 | 0.010509 | 0.10095 | 0.010509 | 0.10095 | -19.8713 | 84.0568 | 0.0052556 | 2.551 | 2.551 | 89.882 | -0.10002 | -0.9946 | 0.10186 | -0.9936 | 0.10176 | -0.9936 | 0.10176 | -0.9936 | 0.10176 | -0.010423 | 174.1527 | 0.030078 | 2.5536 | 2.5538 | 89.3252 | 0.0067098
199 | 870963589.9561 | 0.005 | 2.7362 | 2.7362 | 89.8953 |  | 0.01203 | 0.10811 | 0.012018 | 0.108 | 0.012018 | 0.108 | 0.012018 | 0.108 | -19.2779 | 83.6505 | 0.0052942 | 2.7335 | 2.7335 | 89.889 | -0.10002 | -0.99383 | 0.1091 | -0.99284 | 0.10899 | -0.99284 | 0.10899 | -0.99284 | 0.10899 | -0.010422 | 173.7353 | 0.030087 | 2.7362 | 2.7364 | 89.37 | 0.0058434
200 | 933254300.797 | 0.005 | 2.9319 | 2.9319 | 89.9023 |  | 0.013759 | 0.11564 | 0.013745 | 0.11552 | 0.013745 | 0.11552 | 0.013745 | 0.11552 | -18.6855 | 83.2147 | 0.0053385 | 2.929 | 2.929 | 89.8956 | -0.10002 | -0.99295 | 0.11685 | -0.99196 | 0.11673 | -0.99196 | 0.11673 | -0.99196 | 0.11673 | -0.010421 | 173.2883 | 0.030099 | 2.9319 | 2.9321 | 89.4118 | 0.0050887
201 | 1000000000 | 0.005 | 3.1416 | 3.1416 | 89.9088 |  | 0.015737 | 0.12366 | 0.015721 | 0.12354 | 0.015721 | 0.12354 | 0.015721 | 0.12354 | -18.0942 | 82.7478 | 0.0053894 | 3.1384 | 3.1385 | 89.9016 | -0.10002 | -0.99194 | 0.12514 | -0.99095 | 0.12502 | -0.99095 | 0.12502 | -0.99095 | 0.12502 | -0.010421 | 172.8095 | 0.030111 | 3.1416 | 3.1417 | 89.4509 | 0.0044315

## Info
File: VNA_impedance_v-w01.xls For updates and with comments contact:
Application information and hints:
1., This file uses dot (.) as separator between the integer and fractional parts of numbers. Preferences in your Microsoft Excel options should be set accordingly.
2., Enter your data in the top left area of the Time_harm worksheet. Select the required input parameter (red) and enter your data in the units shown.
3., There is no error checking of most of your entered input data. Physically invalid input (e.g., negative capacitance) may result in meaningless output.
4., For explanatory information, see the worksheet "Reference". For suggested input data to explore, use the Workgroup Scenario toolbar.
5., If Automatic Recalculation is selected in the Tools/Options menu, numerical and graphical outputs are automatically updated after each new entry. To update the added random noise content, press F9.
6., For best viewing, set the current window to display all input output fields and graphs.
7., The calculation part of the worksheets can be viewed and changed if necessary below the input/output areas. To navigate among cells, you can switch back on the Row and Column Headers and Scroll bars in the Tools/Options menu.

## Reference
Equivalent circuits, and expressions used on the worksheets
This spreadsheet illustrates some of the error contributors when a PDN component, such as a single capacitor defined by its C capacitance, R resistance and L inductance, is measured with a Vector Network Analyzer (VNA), based on one-port reflection (S11) versus two-port transmission (S21) measurements.
The equivalent circuit when we measure impedance based on one-port S11 VNA data, is shown below. Note that in this spreadsheet illustration we focus on potential errors introduced by the VNA and therefore the Zconnection discontinuity is assumed to be negligible.
When we measure low impedances, the magnitude of the voltage reflection coefficient becomes close to one, making the above formula, which calculates impedance, very sensitive to any small error in the magnitude or phase of the measured data.
When we measure impedance based on two-port S21 transmission data, the equivalent circuit becomes:
The unknown impedance can be calculated from the complex voltage divider between the 25 Ohm VNA impedance (the two 50-ohm ports in parallel) and the Zx unknown impedance. The Through calibration sets half of the source voltage (Vs/2) to unity. For small impedance values the S21 reading is small, and the sensitivity of calculated impedance value to measurement errors has no 'magnification' as in case of the one-port measurement.
This spreadsheet uses a single C-R-L capacitor model as Device Under Test. Tthe complex impedance of the series-equivalent capacitor model:

## Copyright
This toolset is copyright, Istvan Novak. All rights reserved. You are hereby granted a limited right to run, copy, and reproduce the toolset for your personal, non-commercial purposes only. By using the file(s), you acknowledge not to run, show or disseminate in any way the toolset in full or in part at public or commercial presentations without the express written consent of the author.